## Master_IEEE_HL7_Data_Objects_20
Dialysis Machine Data Standard: IEEE 11073 identifiers and containment
REFID | AlertType | DialysisCommonName | UOM_UCUM | Enum_Values | Enum_Descriptions
PID-3 |  | Patient Identifier | -- | . | .
PID-5 |  | Patient Name | -- | . | .
PV1-19 |  | Visit Number | -- | . | .
MDCX_DEV_HDIALY_MACHINE_MDS
. . . MDC_ATTR_ID_MODEL |  | Dialysis Device Model and Manufacturer | -- | . | .
. . . MDC_ID_MODEL_MANUFACTURER |  | Dialysis Device Manufacturer | -- | . | .
. . . MDC_ID_MODEL_NUMBER |  | Dialysis Device Model | -- | . | .
. . . MDC_ID_PROD_SPEC_SERIAL |  | Dialysis Device Serial Number | -- | . | .
. . . MDC_ID_PROD_SPEC_SW |  | Device Software Version | -- | . | .
. . . PRT-10 PRT-16-20 MDC_ATTR_ID_UDI |  | Dialysis Device UDI | -- | . | .
. MDCX_DEV_HDIALY_VMD
. . MDCX_DEV_HDIALY_MACH_CONFIG_CHAN
. . . MDCX_HDIALY_MACH_TIME |  | Dialysis Device Time | -- | . | .
. . . MDCX_HDIALY_MACH_MODE_DESCRIPTION |  | Machine Mode Description | -- | . | .
. . . MDCX_HDIALY_MACH_MODE_OF_OPERATION |  | Machine Mode Of Operation | -- | PRETX | Pre-Treatment
|  |  |  | TX | Treatment
|  |  |  | POSTTX | Post-Treatment
|  |  |  | DIS | Disinfection
|  |  |  | IDL | Idle
|  |  |  | SVC | Service
. . . MDCX_HDIALY_MACH_PM_TIME_REMAIN |  | Preventive Maintenance Time Remaining | h | . | .
. . . MDCX_HDIALY_MACH_MAINT_TX_REMAIN |  | Preventive Maintenance Treatments Remaining | -- | . | .
. . . MDCX_HDIALY_MACH_BLD_PUMP_ON |  | Treatment Blood Pump On | -- | . | .
. . . MDCX_HDIALY_MACH_TX_FLUID_BYPASS |  | Treatment Fluid Bypass | -- | . | .
. . . MDCX_HDIALY_MACH_TX_MODALITY |  | Treatment Modality | -- | HD | Hemodialysis
|  |  |  | HDF | Hemodiafiltration
|  |  |  | HF | Hemofiltration
|  |  |  | SLED | Sustained low efficiency dialysis
|  |  |  | IUF | Isolated Ultrafiltration
|  |  |  | HP | Hemoperfusion
. . . MDCX_HDIALY_MACH_THERAPY_TIME |  | Elapsed Treatment Time | min | . | .
. . . MDCX_HDIALY_MACH_TIME_REMAIN |  | Remaining Treatment Time | min | . | .
. . MDCX_DEV_HDIALY_ANTICOAG_PUMP_CHAN
. . . MDCX_HDIALY_ANTICOAG_ACCUM_DELIV |  | Accumulated Delivered Anticoagulant | mL | . | .
. . . MDCX_HDIALY_ANTICOAG_BOLUS_VOL |  | Anticoagulant Bolus Volume | mL | . | .
. . . MDCX_HDIALY_ANTICOAG_INFUS_RATE |  | Anticoagulant Infusion Rate | mL/h | . | .
. . . MDCX_HDIALY_ANTICOAG_TYPE |  | Anticoagulant Type | -- | . | .
. . . MDCX_HDIALY_ANTICOAG_MODE |  | Anticoagulation Mode | -- | BOL | Bolus
|  |  |  | CON | Continuous
|  |  |  | BOLCON | Bolus and Continuous
|  |  |  | PRO | Profile
|  |  |  | BOLPRO | Bolus and Profile
|  |  |  | NONE | None
. . . MDCX_HDIALY_ANTICOAG_PUMP_NUMBER |  | Number of Anticoagulant Pumps | -- | . | .
. . . MDCX_HDIALY_ANTICOAG_SYRINGE_BRAND |  | Syringe Name | -- | . | .
. . . MDCX_HDIALY_ANTICOAG_SYRINGE_VOL |  | Syringe Volume | mL | . | .
. . MDCX_DEV_HDIALY_BLOOD_PUMP_CHAN
. . . MDCX_HDIALY_BLD_PUMP_BLOOD_FLOW_RATE |  | Actual Blood Flow Rate | mL/min | . | .
. . . MDCX_HDIALY_BLD_PRESS_ART | phys tech high low thr | Arterial Pressure | mm[Hg] | . | .
. . . MDCX_HDIALY_BLD_PUMP_BLOOD_FLOW_RATE_AVG |  | Average Blood Flow Rate | mL/min | . | .
. . . MDCX_HDIALY_BLD_PUMP_BLOOD_FLOW_RATE_AVG_SETTING |  | Blood Flow Rate Setting | mL/min | . | .
. . . MDCX_HDIALY_BLD_PUMP_MODE |  | Blood Pump Mode | -- | 2N | Double Needle
|  |  |  | 1N1P | Single Needle/Single Pump
|  |  |  | 1N2P | Single Needle/Double Pump
. . . MDCX_HDIALY_BLD_PUMP_TUBING_SIZE |  | Blood Tubing Size | -- | . | .
. . . MDCX_HDIALY_BLD_PUMP_PAT_BODY_TEMP_BY_BLOOD |  | Body Temperature | Cel | . | .
. . . MDCX_HDIALY_BLD_PUMP_CHANGE_IN_ENERGY |  | Change In Energy | kJ/h | . | .
. . . MDCX_HDIALY_BLD_PUMP_PRESS_ART_POST_PUMP | tech high low thr | Post Pump Arterial Pressure | mm[Hg] | . | .
. . . MDCX_HDIALY_BLD_PUMP_PRIMING_VOL |  | Priming Volume | mL | . | .
. . . MDCX_HDIALY_BLD_PUMP_SINGLE_NEEDLE_PRESS | tech high low thr | Single Needle Switching Pressure | mm[Hg] | . | .
. . . MDCX_HDIALY_BLD_PUMP_STROKE_VOL |  | Stroke Volume | mL | . | .
. . . MDCX_HDIALY_BLD_PUMP_BLOOD_PROCESSED_TOTAL |  | Total Blood Processed | L | . | .
. . . MDCX_HDIALY_BLD_PUMP_PRESS_VEN | phys tech high low thr | Venous Pressure | mm[Hg] | . | .
. . . MDCX_HDIALY_BLD_PUMP_BLOOD_TEMP_VEN |  | Venous Temperature | Cel | . | .
. . MDCX_DEV_HDIALY_FLUID_CHAN
. . . MDCX_HDIALY_BICARB_COND |  | Actual Bicarbonate Conductivity | mS/cm | . | .
. . . MDCX_HDIALY_BICARB_COND_SETTING |  | Bicarbonate Conductivity Setting | mS/cm | . | .
. . . MDCX_HDIALY_DIALYSATE_COND | tech high low | Actual Conductivity | mS/cm | . | .
. . . MDCX_HDIALY_DIALYSATE_COND_SETTING |  | Dialysate Conductivity Setting | mS/cm | . | .
. . . MDCX_HDIALY_DIALYSATE_FLOW_RATE | tech low | Actual Dialysate Flow Rate | mL/min | . | .
. . . MDCX_HDIALY_DIALYSATE_FLOW_RATE_SETTING |  | Dialysate Flow Rate Setting | mL/min | . | .
. . . MDCX_HDIALY_DIALYSATE_FLOW_MODE |  | Dialysate Flow Mode | -- | CONST | Constant
|  |  |  | AUTO | Auto
|  |  |  | PRO | Profile
|  |  |  | STBY | Standby
. . . MDCX_HDIALY_DIALYSATE_TEMP | tech high low | Actual Dialysate Temperature | Cel | . | .
. . . MDCX_HDIALY_DIALYSATE_TEMP_SETTING |  | Dialysate Temperature Setting | Cel | . | .
. . . MDCX_HDIALY_DIALYSATE_AMMONIA |  | Ammonia Concentration | [ppm] | . | .
. . . MDCX_HDIALY_DIALYSATE_FLOW_RATE_AVG |  | Average Dialysate Flow Rate | mL/min | . | .
. . . MDCX_HDIALY_CONC_HCO3_SETTING |  | Bicarbonate Concentration Setting | mmol/L | . | .
. . . MDCX_HDIALY_CONC_HCO3_MODE |  | Bicarbonate Mode | -- | CONST | Constant
|  |  |  | PRO | Profile
|  |  |  | NA | Not Applicable
. . . MDCX_HDIALY_DIALYSATE_NAME |  | Dialysate Name | -- | . | .
. . . MDCX_HDIALY_DIALYSATE_CONC_ACETATE |  | Dialysate Acetate Concentration | mmol/L | . | .
. . . MDCX_HDIALY_DIALYSATE_CONC_HCO3 |  | Dialysate Bicarbonate Concentration | mmol/L | . | .
. . . MDCX_HDIALY_DIALYSATE_CONC_CHLORIDE |  | Dialysate Chloride Concentration | mmol/L | . | .
. . . MDCX_HDIALY_DIALYSATE_CONC_MG |  | Dialysate Magnesium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_DIALYSATE_CONC_CA |  | Dialysate Calcium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_DIALYSATE_CONC_CITRATE |  | Dialysate Citrate Concentration | mmol/L | . | .
. . . MDCX_HDIALY_DIALYSATE_CONC_GLU |  | Dialysate Glucose Concentration | mmol/L | . | .
. . . MDCX_HDIALY_DIALYSATE_CONC_K |  | Dialysate Potassium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_DIALYSATE_CONC_NA |  | Dialysate Sodium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_DIALYSATE_CONC_NA_SETTING |  | Dialysate Sodium concentration Setting | mmol/L | . | .
. . . MDCX_HDIALY_DIALYSATE_CONC_NA_MODE |  | Dialysate Sodium Mode | -- | CONST | Constant
|  |  |  | PRO | Profile
. . . MDCX_HDIALY_DIALYSATE_CONC_PH |  | pH | [pH] | . | .
. . . MDCX_HDIALY_DIALYSATE_VOL_DELIV |  | Total Dialysate Volume | L | . | .
. . MDCX_DEV_HDIALY_FILTER_CHAN
. . . MDCX_HDIALY_FILTER_NAME |  | Dialyzer Name | -- | . | .
. . . MDCX_HDIALY_FILTER_TRANSMEMBRANE_PRESS | tech high low | Actual Transmembrane Pressure | mm[Hg] | . | .
. . MDCX_DEV_HDIALY_CONVECTIVE_CHAN
. . . MDCX_HDIALY_RF_DILUTION_LOCATION |  | Replacement Fluid Dilution Location | -- | PREF | Pre-Filter
|  |  |  | POSTF | Post-Filter
|  |  |  | PREF-POSTF | Pre-Filter/Post-Filter
|  |  |  | PREP-POSTF | Pre-Pump/Post-Filter
. . . MDCX_HDIALY_RF_POST_FILTER_NAME |  | Post-Filter RF Solution Name | -- | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_CONC_ACETATE |  | Post-Filter RF Acetate Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_CONC_HCO3 |  | Post-Filter RF Bicarbonate Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_CONC_CHLORIDE |  | Post-Filter RF Chloride Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_CONC_MG |  | Post-Filter RF Magnesium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_CONC_CA |  | Post-Filter RF Calcium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_CONC_CITRATE |  | Post-Filter RF Citrate Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_CONC_GLU |  | Post-Filter RF Glucose Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_CONC_K |  | Post-Filter RF Potassium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_CONC_NA |  | Post-Filter RF Sodium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_FLOW_RATE |  | Actual Post-Filter Replacement Fluid Flow Rate | mL/min | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_FLOW_RATE_SETTING |  | Post-Filter Replacement Fluid Flow Rate Setting | mL/min | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_FLOW_RATE_AVG |  | Average Post-Filter Replacement Fluid Flow Rate | mL/min | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_TEMP |  | Actual Post-Filter Replacement Fluid Temperature | Cel | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_TEMP_SETTING |  | Post-Filter Replacement Fluid Temperature Setting | Cel | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_VOL |  | Total Post-Filter Replacement Fluid Volume | L | . | .
. . . MDCX_HDIALY_RF_POST_FILTER_VOL_SETTING |  | Post-Filter Replacement Fluid Volume setting | L | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_NAME |  | Pre-Filter RF Solution Name | -- | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_CONC_ACETATE |  | Pre-Filter RF Acetate Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_CONC_HCO3 |  | Pre-Filter RF Bicarbonate Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_CONC_CHLORIDE |  | Pre-Filter RF Chloride Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_CONC_MG |  | Pre-Filter RF Magnesium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_CONC_CA |  | Pre-Filter RF Calcium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_CONC_CITRATE |  | Pre-Filter RF Citrate Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_CONC_GLU |  | Pre-Filter RF Glucose Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_CONC_K |  | Pre-Filter RF Potassium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_CONC_NA |  | Pre-Filter RF Sodium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_FLOW_RATE |  | Actual Pre-Filter Replacement Fluid Flow Rate | mL/min | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_FLOW_RATE_SETTING |  | Pre-Filter Replacement Fluid Flow Rate Setting | mL/min | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_FLOW_RATE_AVG |  | Average Pre-Filter Replacement Fluid Flow Rate | mL/min | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_TEMP |  | Actual Pre-Filter Replacement Fluid Temperature | Cel | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_TEMP_SETTING |  | Pre-Filter Replacement Fluid Temperature Setting | Cel | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_VOL |  | Total Pre-Filter Replacement Fluid Volume | L | . | .
. . . MDCX_HDIALY_RF_PRE_FILTER_VOL_SETTING |  | Pre-Filter Replacement Fluid Volume Setting | L | . | .
. . . MDCX_HDIALY_RF_CONV_CLEARANCE |  | Convective Clearance - Urea | mL/min | . | .
. . . MDCX_HDIALY_RF_PRE_POST_FLOW_RATIO_SETTING |  | Pre-Post Replacement Fluid Flow Rate Ratio Setting | -- | . | .
. . . MDCX_HDIALY_RF_BOLUS_RATE | tech high low | Replacement Fluid Bolus Rate | mL/min | . | .
. . . MDCX_HDIALY_RF_BOLUS_VOL_SETTING |  | Replacement Fluid Bolus Volume Setting | mL | . | .
. . . MDCX_HDIALY_RF_BOLUS_VOL_DELIVERED | tech high low | Replacement Fluid Bolus Volume Delivered | mL | . | .
. . . MDCX_HDIALY_RF_FLOW_MODE |  | Replacement Fluid Flow Mode | -- | CONST | Constant
|  |  |  | PRO | Profile
|  |  |  | NONE | None or Not Applicable
|  |  |  | BOLUS | Bolus
|  |  |  | AUTO | Automatic
. . . MDCX_HDIALY_RF_SOURCE |  | Replacement Fluid Source | -- | BAG | Bag
|  |  |  | ONLINE | Online
. . MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN
. . MDCX_DEV_HDIALY_THERAPY_OUTCOMES_CHAN
. . . MDCX_HDIALY_THERAPY_MASS_TRF_AREA_COEFF |  | KoA | -- | . | .
. . . MDCX_HDIALY_THERAPY_ACCESS_FLOW |  | Access Flow | mL/min | . | .
. . . MDCX_HDIALY_THERAPY_RATIO_EKT_OVER_V_DELIVERED |  | Delivered Equilibrated Kt/V | % | . | .
. . . MDCX_HDIALY_THERAPY_KT_DELIVERED |  | Delivered Kt | L | . | .
. . . MDCX_HDIALY_THERAPY_RATIO_SPKT_OVER_V_DELIVERED |  | Delivered Single Pool Kt/V | % | . | .
. . . MDCX_HDIALY_THERAPY_RATIO_KT_OVER_V_GOAL |  | Kt/V Goal | % | . | .
. . . MDCX_HDIALY_THERAPY_MEAN_UREA_CLEARANCE |  | Mean Urea Clearance | -- | . | .
. . . MDCX_HDIALY_THERAPY_BODY_START_WT |  | Patient Start Weight | kg | . | .
. . . MDCX_HDIALY_THERAPY_PCT_RECIRC |  | Percent Recirculation | % | . | .
. . . MDCX_HDIALY_THERAPY_PLASMA_NA_CONC |  | Plasma Sodium Concentration | mmol/L | . | .
. . . MDCX_HDIALY_THERAPY_RATIO_SPKT_OVER_V_PROJECTED |  | Projected Single Pool Kt/V | % | . | .
. . . MDCX_HDIALY_THERAPY_COMPLETE_METHOD |  | Treatment Completion Method | -- | CT | Clock Time
|  |  |  | AT | Actual Treatment Time
|  |  |  | UF | UF Removed
|  |  |  | KTV | Kt/V
. . . MDCX_HDIALY_THERAPY_UREA_CLEARANCE |  | Urea Clearance | -- | . | .
. . . MDCX_HDIALY_THERAPY_BODY_END_WT |  | Patient End Weight | kg | . | .
. . MDCX_DEV_HDIALY_UF_CHAN
. . . MDCX_HDIALY_UF_TIME_TO_TARGET |  | Remaining Ultrafiltration Time | min | . | .
. . . MDCX_HDIALY_UF_TARGET_VOL_TO_REMOVE |  | Target Volume To Remove | mL | . | .
. . . MDCX_HDIALY_UF_ACTUAL_REMOVED_VOL |  | Total Fluid Volume Removed | mL | . | .
. . . MDCX_HDIALY_UF_RATE_MAX_NON_ALM_SETTING |  | Ultrafiltration Flow Rate Maximum Limit | mL/h | . | .
. . . MDCX_HDIALY_UF_MODE |  | Ultrafiltration Mode | -- | CIUF | Constant Isolated UF
|  |  |  | PIUF | Profile Isolated UF
|  |  |  | CUF | Constant UF
|  |  |  | PUF | Profile UF
|  |  |  | NONE | None
. . . MDCX_HDIALY_UF_RATE |  | Ultrafiltration Rate | mL/h | . | .
. MDC_DEV_PRESS_BLD_NONINV_VMD
. . MDC_DEV_PRESS_BLD_NONINV_CHAN
. . . MDC_ATTR_TIME_PD_MSMT |  | Blood Pressure Measurement Interval Setting | min | . | .
. . . MDCX_ATTR_BODY_POSN |  | Blood Pressure Measurement Position | -- | SITTING | Sitting
|  |  |  | STANDIND | Standing
|  |  |  | SUPINE | Supine
. . . MDC_PRESS_BLD_NONINV_DIA | phys high low thr | Diastolic Pressure | mm[Hg] | . | .
. . . MDC_PULS_RATE_NON_INV | phys high low thr | Heart Rate | {beats}/min | . | .
. . . MDC_PRESS_BLD_NONINV_MEAN |  | Mean Arterial Pressure | mm[Hg] | . | .
. . . MDC_PRESS_BLD_NONINV_SYS | phys high low thr | Systolic Pressure | mm[Hg] | . | .
. MDC_DEV_ANALY_SAT_O2_VMD
. . MDC_DEV_ANALY_SAT_O2_CHAN
. . . MDC_PULS_OXIM_SAT_O2 | phys high low thr | SpO2 Oxygen Saturation | % | . | .
. . . MDC_PULS_OXIM_PULS_RATE | phys high low thr | SpO2 Pulse Rate | {beats}/min | . | .
. MDCX_DEV_BLOOD_CHEM_VMD
. . MDCX_DEV_BLOOD_CHEM_CHAN
. . . MDCX_HDIALY_PLASMA_VOL_MARKER |  | Plasma Volume Marker | -- | None | None – No selection made at time of reading/reporting
|  |  |  | SYMP | Symptomatic – Noting a symptom for continue monitoring.
|  |  |  | INTER | Intervention – Action was taken or a machine setting was changed.
. . . MDCX_HDIALY_PLASMA_VOL_PROFILE |  | Plasma Volume Profile | -- | FLAT | Change less than or equal to 3% per hour
|  |  |  | GRAD | Change greater than 3% and less than or equal to 6.5% per hour
|  |  |  | STEEP | Change greater than 6.5% per hour.
. . . MDC_CONC_HCT | phys low high thr | Hematocrit | %{vol} | . | .
. . . MDC_CONC_HB | phys low high thr | Hemoglobin | g/dL | . | .
. . . MDCX_HDIALY_REL_BLOOD_VOL | phys low thr | Relative Blood Volume | % | . | .
. . . MDC_SAT_O2 | phys low high thr | SatO2 Oxygen Saturation | % | . | .
IEEE 11073 Dialysis event and alert source and event identifiers
source | event | AlertType | DialysisCommonName
MDCX_DEV_HDIALY_MACHINE_MDS
. MDCX_DEV_HDIALY_VMD
. . MDCX_DEV_HDIALY_MACH_CONFIG_CHAN
. . MDCX_DEV_HDIALY_ANTICOAG_PUMP_CHAN
MDCX_DEV_HDIALY_ANTICOAG_PUMP_CHAN | MDCX_EVT_HDIALY_ANTICOAG_STOP | tech | Anticoagulant Pump Stop
MDCX_DEV_HDIALY_ANTICOAG_PUMP_CHAN | MDCX_EVT_HDIALY_ANTICOAG_SYRINGE_EMPTY | tech | Syringe Empty
MDCX_DEV_HDIALY_ANTICOAG_PUMP_CHAN | MDCX_EVT_HDIALY_ANTICOAG_SYRINGE_SIZE | tech | Wrong Syringe Size
. . MDCX_DEV_HDIALY_BLOOD_PUMP_CHAN
MDCX_HDIALY_BLD_PRESS_ART | MDC_EVT_LO, MDC_EVT_HI | phys tech high low thr | Arterial Pressure
MDCX_DEV_HDIALY_BLOOD_PUMP_CHAN | MDCX_EVT_HDIALY_BLD_PUMP_STOP | tech | Blood Pump Stop
MDCX_HDIALY_BLD_PUMP_PRESS_ART_POST_PUMP | MDC_EVT_LO, MDC_EVT_HI | tech high low thr | Post Pump Arterial Pressure
MDCX_HDIALY_BLD_PUMP_SINGLE_NEEDLE_PRESS | MDC_EVT_LO, MDC_EVT_HI | tech high low thr | Single Needle Switching Pressure
MDCX_HDIALY_BLD_PUMP_PRESS_VEN | MDC_EVT_LO, MDC_EVT_HI | phys tech high low thr | Venous Pressure
. . MDCX_DEV_HDIALY_FLUID_CHAN
MDCX_HDIALY_DIALYSATE_COND | MDC_EVT_LO, MDC_EVT_HI | tech high low | Actual Conductivity
MDCX_HDIALY_DIALYSATE_FLOW_RATE | MDC_EVT_LO | tech low | Actual Dialysate Flow Rate
MDCX_HDIALY_DIALYSATE_TEMP | MDC_EVT_LO, MDC_EVT_HI | tech high low | Actual Dialysate Temperature
MDCX_DEV_HDIALY_FLUID_CHAN | MDCX_EVT_HDIALY_BLOOD_LEAK | tech | Blood Leak
. . MDCX_DEV_HDIALY_FILTER_CHAN
MDCX_HDIALY_FILTER_TRANSMEMBRANE_PRESS | MDC_EVT_LO, MDC_EVT_HI | tech high low | Actual Transmembrane Pressure
. . MDCX_DEV_HDIALY_CONVECTIVE_CHAN
MDCX_DEV_HDIALY_CONVECTIVE_CHAN | MDCX_EVT_HDIALY_RF_EXCESS_DELIV | tech-high | Excessive Replacement Fluid Bolus Delivered
MDCX_DEV_HDIALY_CONVECTIVE_CHAN | MDCX_EVT_HDIALY_RF_INSUFF_DELIV | tech-low | Insufficient Replacement Fluid Bolus Delivered
MDCX_HDIALY_RF_BOLUS_RATE | MDC_EVT_LO, MDC_EVT_HI | tech high low | Replacement Fluid Bolus Rate
MDCX_HDIALY_RF_BOLUS_VOL_DELIVERED | MDC_EVT_LO, MDC_EVT_HI | tech high low | Replacement Fluid Bolus Volume Delivered
. . MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN
MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_ART_AIR_DETECT | tech | Arterial Air Detector
MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_DIALYSATE_AIR_DETECT | tech | Dialysate Air Detector
MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_DIALYSATE_COMPOSITION | tech | Dialysis Composition
MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_SYSTEM_GENERAL | tech | General System
MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_SELF_TEST | tech | Self-Test
MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_VEN_ACCESS | tech | Venous Access
MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_VEN_AIR_DETECT | tech | Venous Air Detector
MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_WETNESS_DETECT_ALERT | tech | Wetness Detector
MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_WETNESS_DETECT_ERROR | tech | Wetness Detector Error
. . MDCX_DEV_HDIALY_THERAPY_OUTCOMES_CHAN
MDCX_DEV_HDIALY_THERAPY_OUTCOMES_CHAN | MDCX_EVT_HDIALY_THERAPY_PAT_TX_GENERAL | phys | General Patient Treatment
MDCX_DEV_HDIALY_THERAPY_OUTCOMES_CHAN | MDCX_EVT_HDIALY_THERAPY_TX_END_TIME | tech | Treatment Time Ended
. . MDCX_DEV_HDIALY_UF_CHAN
MDCX_DEV_HDIALY_UF_CHAN | MDCX_EVT_HDIALY_UF_LO | tech-low | Insufficient Ultrafiltration
MDCX_DEV_HDIALY_UF_CHAN | MDCX_EVT_HDIALY_UF_NEG | tech-low | Negative Ultrafiltration
MDCX_DEV_HDIALY_UF_CHAN | MDCX_EVT_HDIALY_UF_GOAL_MET | tech | Ultrafiltration Goal Met
MDCX_DEV_HDIALY_UF_CHAN | MDCX_EVT_HDIALY_UF_RATE_RANGE | tech-range | Ultrafiltration Rate Out of Range
. MDC_DEV_PRESS_BLD_NONINV_VMD
. . MDC_DEV_PRESS_BLD_NONINV_CHAN
MDC_PRESS_BLD_NONINV_DIA | MDC_EVT_LO, MDC_EVT_HI | phys high low thr | Diastolic Pressure
MDC_PULS_RATE_NON_INV | MDC_EVT_LO, MDC_EVT_HI | phys high low thr | Heart Rate
MDC_PRESS_BLD_NONINV_SYS | MDC_EVT_LO, MDC_EVT_HI | phys high low thr | Systolic Pressure
. MDC_DEV_ANALY_SAT_O2_VMD
. . MDC_DEV_ANALY_SAT_O2_CHAN
MDC_PULS_OXIM_SAT_O2 | MDC_EVT_LO, MDC_EVT_HI | phys high low thr | SpO2 Oxygen Saturation
MDC_PULS_OXIM_PULS_RATE | MDC_EVT_LO, MDC_EVT_HI | phys high low thr | SpO2 Pulse Rate
MDC_DEV_ANALY_SAT_O2_CHAN | MDCX_EVT_PULS_OXIM_ERROR | tech | Pulse Oximeter Error
. MDCX_DEV_BLOOD_CHEM_VMD
. . MDCX_DEV_BLOOD_CHEM_CHAN
MDC_CONC_HCT | MDC_EVT_LO, MDC_EVT_HI | phys low high thr | Hematocrit
MDC_CONC_HB | MDC_EVT_LO, MDC_EVT_HI | phys low high thr | Hemoglobin
MDCX_HDIALY_REL_BLOOD_VOL | MDC_EVT_LO | phys low thr | Relative Blood Volume
MDC_SAT_O2 | MDC_EVT_LO, MDC_EVT_HI | phys low high thr | SatO2 Oxygen Saturation
NOTES:
Numeric limit events and alerts are identified by the source::event pair { numeric observation REFID :: MDC_EVT_LO ¦ MDC_EVT_HI }
Most technical and other 'named' events and alerts are identified by { MDS, VMD, CHAN or numeric observation REFID :: MDC_EVT_HEMODYALYSIS or other MDC_EVT code }
This document is pre-ballot draft material for proposed IEEE Standard, and as such, it is subject to change. USE AT YOUR OWN RISK!
Because this is based on an unapproved draft content, this document must not be utilized for any conformance/compliance purposes.