## DataAndAlerts
REFID | SysName | SysNameBase | SysName1stDesc | SysName2ndDesc | SysName3rdDesc | AlertType | Group | DialysisDeviceVMD | CommonTerm | Description | Phase | Message | Temporal | DataType | Format | UOM | UOM_UCUM | UOM_MDC | OneAnyOf | MOC-Status | Enum_Tbl | Enum_Values | Enum_Descriptions | Range | Example | gp | s | Disc | PART | pCODE10 | CODE10 | CF_CODE10
PID-3 | [6 empty] | HEMODIALYSIS | Patient Identifier | Patient Identifier | Unique alphanumeric value assigned to the person receiving dialysis treatment | Intradialytic | Identifier | All | String | Alphanumeric | N/A | -- | -- | 1..1 | M |  | -- | -- | N/A | N/A
PID-5 | [6 empty] | HEMODIALYSIS | Patient Identifier | Patient Name | The last and first name of the person receiving dialysis treatment | Intradialytic | Identifier | All | String | Alphanumeric | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | N/A
PV1-19 | [6 empty] | HEMODIALYSIS | Patient Identifier | Visit Number | A value assigned to define treatment schedule for patient | Intradialytic | Identifier | All | String | Alphanumeric | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | N/A
Profile Parameters | [6 empty] | HEMODIALYSIS
. . . . MDC_HDIALY_PROFILE_TYPE | [6 empty] | HEMODIALYSIS | N/A | Profile Type | Contains the type of the profile. This will be VENDOR, CONSTANT, LINEAR, EXPONENTIAL, or STEP. | All | Parameter | Episodic | Enumeration | N/A | N/A
. . . . MDC_HDIALY_PROFILE_VALUE | [6 empty] | HEMODIALYSIS | N/A | Profile Values | Contains the control points of the profile. The control points are a numeric array. For a constant profile there is only one point. For a linear or exponential profile there are two points, one for the starting value and one for the ending value. For step profile there is one value for each step. | All | Parameter | Episodic | Array | Numeric | same as parameter being profiled.
. . . . MDC_HDIALY_PROFILE_TIME | [6 empty] | HEMODIALYSIS | N/A | Profile Times | This parameter is optional. If it is not present, the control points are assumed to be equally space across the total treatment time. If it is present, it is an array of numerical values that are the treatment times at which a control point takes effect. | All | Parameter | Episodic | Array | Numeric | minutes
. . . . MDC_HDIALY_PROFILE_EXP_HALF_TIME | [6 empty] | HEMODIALYSIS | N/A | Profile Half Time | This parameter is optional and only applies to the Exponential Profile. It specifies the time at which 50% of the change has been complete. | All | Parameter | Episodic | Numeric | XX | minutes
. . . . MDC_HDIALY_PROFILE_NAME | [6 empty] | HEMODIALYSIS | N/A | Profile Parameters | This parameter is optional and only applies to the Vendor Profile. It uniquely identifies the profile. | All | Parameter | Episodic | String |  | N/A
. . . . MDC_HDIALY_PROFILE_PARAMETERS | [6 empty] | HEMODIALYSIS | N/A | Profile Parameters | Vendor specific parameters for a vendor specific profile. | All | Parameter | Episodic | Array | Numeric | vendor specific
MDC_DEV_HDIALY_MACHINE_MDS | System ¦ ¦ ¦ Dialysis | System |  |  | Dialysis |  |  |  | Hemodialysis machine | A device or system used to treat renal failure delivered by removing toxins from the blood through an artificial kidney, called a dialyzer. | [8 empty] | 1..1 | M | [5 empty] | 3 |  | MVC (1) | 1 |  | 5393 | 70929
. . . MDC_ATTR_ID_MODEL | System (string) ¦ Identifier ¦ Model Description ¦ Device | System (string) | Identifier | Model Description | Device |  | HEMODIALYSIS | Dialysis Device | Dialysis Device Model and Manufacturer | String containing device model and manufacturer's name | All | Identifier | All | String | Alphanumeric | N/A | -- | -- | 1..1 | M |  | -- | -- | N/A | NxStage System One | P |  | 1 | 1 | 2344 | 2344 | 67880
. . . MDC_ID_MODEL_MANUFACTURER | System (string) ¦ Identifier ¦ Manufacturer ¦ Device | System (string) | Identifier | Manufacturer | Device |  | HEMODIALYSIS | Dialysis Device | Dialysis Device Manufacturer | String containing device manufacturer's name (or use MDC_ATTR_ID_MODEL) | All | Identifier | All | String | Alphanumeric | N/A | -- | -- | 1..1 | M |  | -- | -- | N/A | NxStage | P |  | 1 | 8 | 7682 | 7682 | 531970
. . . MDC_ID_MODEL_NUMBER | System (string) ¦ Identifier ¦ Model Number ¦ Device | System (string) | Identifier | Model Number | Device |  | HEMODIALYSIS | Dialysis Device | Dialysis Device Model | String containing device manufacturer's model identifier for the device (or use MDC_ATTR_ID_MODEL) | All | Identifier | All | String | Alphanumeric | N/A | -- | -- | 1..1 | M |  | -- | -- | N/A | System One | P |  | 1 | 8 | 7681 | 7681 | 531969
. . . MDC_ID_PROD_SPEC_SERIAL | System (string) ¦ Identifier ¦ Serial Number ¦ Device | System (string) | Identifier | Serial Number | Device |  | HEMODIALYSIS | Dialysis Device | Dialysis Device Serial Number | String containing the device’s serial number | All | Identifier | All | String | Alphanumeric | N/A | -- | -- | 1..1 | M |  | -- | -- | N/A | AS1000478 | P |  | 1 | 8 | 7684 | 7684 | 531972
. . . MDC_ID_PROD_SPEC_SW | System (string) ¦ Identifier ¦ Software Version ¦ Device | System (string) | Identifier | Software Version | Device |  | HEMODIALYSIS | Dialysis Device | Device Software Version | String defining the version of the software on the machine | All | Identifier | All | String | Alphanumeric | N/A | -- | -- | 1..1 | M |  | -- | -- | N/A | 1.2.3.4 | P |  | 1 | 8 | 7687 | 7687 | 531975
. . . MDC_ATTR_ID_UDI | System (string) ¦ Identifier, Unique ¦ Universal identifier, UDI ¦ Device | System (string) | Identifier, Unique | Universal identifier, UDI | Device |  | HEMODIALYSIS | Dialysis Device | Dialysis Device UDI | String containing the device’s FDA required Universal Device Identifier [Note: MDC_ATTR_ID_UDI contains three components: UdiAuthority, UdiIssuer, UdiLabel to support non-FDA entities, but currently there is no approved way of encoding this using HL7 V2; PRT-10 and PRT-16-20 is used instead; UdiLabel = PRT-10.] | All | Identifier | All | String | Alphanumeric | N/A | -- | -- | 1..1 | O |  | -- | -- | N/A | +M535NX10003A0/$$+735241/16D20180305J | P |  | 1 | 1 | 2380 | 2380 | 67916
. MDC_DEV_HDIALY_VMD | Device ¦ ¦ ¦ Dialysis | Device |  |  | Dialysis |  |  |  | Hemodialysis subsystem | A device or subsystem used to treat renal failure delivered by removing toxins from the blood through an artificial kidney, called a dialyzer. | [8 empty] | 1..1 | M | [5 empty] | 3 |  | MVC (2) | 1 |  | 5398 | 70934
. . MDC_DEV_HDIALY_MACH_CONFIG_CHAN | Subsystem ¦ ¦ Machine Configuration ¦ Dialysis | Subsystem |  | Machine Configuration | Dialysis |  |  |  | Hemodialysis configuration | The principal dialysis machine operational modes, treatment modality, settings and other information about the dialysis machine. | [8 empty] | 1..1 | M | [5 empty] | 3 |  | MVC (3) | 1 |  | 5403 | 70939
. . . MDC_HDIALY_MACH_TIME | Time ¦ Machine Time ¦ Machine Configuration ¦ Dialysis | Time | Machine Time | Machine Configuration | Dialysis |  | HEMODIALYSIS | Dialysis Device | Dialysis Device Time | Date and time as recorded on the dialysis device's internal clock with offset to UTC (based on location and DST) | All | Identifier | All | Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | -- | 1..1 | M |  | -- | -- | N/A | 20190421130045+/-0000 | 3 |  | 1 | 2 |  | 27520 | 158592
. . . MDC_HDIALY_MACH_MODE_DESCRIPTION | Mode (string) ¦ Manufacturer-specific mode or system description ¦ Machine Configuration ¦ Dialysis | Mode (string) | Manufacturer-specific mode or system description | Machine Configuration | Dialysis |  | HEMODIALYSIS | Dialysis Device | Machine Mode Description | Manufacturer-specific description of the machine mode of operation | All | Parameter | Episodic | String | Alphanumeric | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | NxStage System One | 3 |  | 1 | 2 |  | 27521 | 158593
. . . . MDC_HDIALY_MACH_MODE_OF_OPERATION | Mode (enumeration) ¦ Operational mode ¦ Machine Configuration ¦ Dialysis | Mode (enumeration) | Operational mode | Machine Configuration | Dialysis |  | HEMODIALYSIS | Dialysis Device | Machine Mode Of Operation | Process the machine is currently performing | All | Parameter | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 1..1 | M | _TBL_01 – Mode of Operation | PRETX TX POSTTX DIS IDL SVC | Pre-Treatment – Preparing for dialysis but no blood in extracorporeal circuit Treatment – Blood in the extracorporeal circuit Post-Treatment – After dialysis but no longer any blood in extracorporeal circuit Disinfection or rinse of dialysis fluid path. Idle Service mode | N/A | TX | 3 |  | 1 | 2 |  | 27522 | 158594
. . . MDC_TIME_PD_MAINTENANCE_TO_NEXT_SERVICE | Duration ¦ Time Remaining, until next preventative maintenance ¦ Machine Configuration ¦ Dialysis | Duration | Time Remaining, until next preventative maintenance | Machine Configuration | Dialysis |  | HEMODIALYSIS | Dialysis Device | Preventive Maintenance Time Remaining | Hours of operation remaining until next preventative maintenance is required | All | Status | Episodic | Numeric | XXXX | hours | h | MDC_DIM_HR | 0..1 | O |  | -- | -- | 0-9999 | 12 | 2 |  | 1 | 2 | 53123 | 53123 | 184195
. . . MDC_MAINTENANCE_NEXT_SERVICE_DATE | Date ¦ Due Date, Preventative Maintenance ¦ Equipment Management ¦ Machine Configuration ¦ Dialysis | Date | Due Date, Preventative Maintenance ¦ Equipment Management | Machine Configuration | Dialysis |  | HEMODIALYSIS | Dialysis Device | Preventative Maintenance Due Date | The date on which preventative maintenance is due. | All | Status | Episodic | Date | YYYYMMDD | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | 20210306 | 2 |  | 1 | 2 | 53127 | 53127 | 184199
. . . MDC_HDIALY_MACH_MAINT_TX_REMAIN | Number ¦ Treatments Remaining, Preventative Maintenance ¦ Machine Configuration ¦ Dialysis | Number | Treatments Remaining, Preventative Maintenance | Machine Configuration | Dialysis |  | HEMODIALYSIS | Dialysis Device | Preventive Maintenance Treatments Remaining | Number of treatments remaining until next preventative maintenance is required | All | Status | Episodic | Numeric | XXXX | N/A | -- | -- | 0..1 | O |  | -- | -- | 0-9999 | 45 | 3 |  | 1 | 2 |  | 27523 | 158595
. . . MDC_HDIALY_MACH_BLD_PUMP_ON | Status (boolean) ¦ Treatment blood pump, On/Off ¦ Machine Configuration ¦ Dialysis | Status (boolean) | Treatment blood pump, On/Off | Machine Configuration | Dialysis |  | HEMODIALYSIS | Dialysis Device | Treatment Blood Pump On | Status indicating if blood pump is running | Intradialytic | Parameter | Episodic | Boolean | T / F | N/A | -- | -- | 1..1 | M |  | -- | -- | N/A | T | 3 |  | 1 | 2 |  | 27524 | 158596
. . . MDC_HDIALY_MACH_TX_FLUID_BYPASS | Status (boolean) ¦ Treatment fluid flow, Bypass On/Off (fluid is NOT flowing through dialyzer) ¦ Machine Configuration ¦ Dialysis | Status (boolean) | Treatment fluid flow, Bypass On/Off (fluid is NOT flowing through dialyzer) | Machine Configuration | Dialysis |  | HEMODIALYSIS | Dialysis Device | Treatment Fluid Bypass | Fluid pathway state to indicate if fluid is NOT flowing through dialyzer | Intradialytic | Parameter | Episodic | Boolean | T / F | N/A | -- | -- | 1..1 | M |  | -- | -- | N/A | F | 3 |  | 1 | 2 |  | 27525 | 158597
. . . MDC_HDIALY_MACH_TX_MODALITY | Treatment (enumeration) ¦ Treatment Modality ¦ Machine Configuration ¦ Dialysis | Treatment (enumeration) | Treatment Modality | Machine Configuration | Dialysis |  | HEMODIALYSIS | Dialysis Device | Treatment Modality | The way a patient receives dialysis | All | Parameter | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 1..1 | M | _TBL_02 – Treatment Modality | HD HDF HF SLED IUF HP | Hemodialysis – dialysis with diffusive transport of solutes. Hemodiafiltration – dialysis with both diffusive and convective transport of solutes Hemofiltration – dialysis with convective transport of solutes. Sustained low efficiency dialysis. Isolated Ultrafiltration – Removal of plasma water without dialysis Hemoperfusion – Removal of solutes from blood by adsorption. | N/A | HDF | 3 |  | 1 | 2 |  | 27526 | 158598
. . . MDC_HDIALY_MACH_THERAPY_TIME | Duration ¦ Accumulated treatment time, Patient actively treated (excluding alarms, recirculating, rinsing back) ¦ Machine Configuration ¦ Dialysis | Duration | Accumulated treatment time, Patient actively treated (excluding alarms, recirculating, rinsing back) | Machine Configuration | Dialysis |  | HEMODIALYSIS | Dialysis Device State | Elapsed Treatment Time | Accumulated time that the patient has been actively treated. This does not include time spend in alarms, recirculating or rinsing back. | Intradialytic | Status | Periodic | Numeric | XXXX | minutes | min | MDC_DIM_MIN | 1..1 | M |  | -- | -- | 0-9999 | 180 | 3 |  | MMM | 2 |  | 27648 | 158720
. . . MDC_HDIALY_MACH_TIME_REMAIN | Duration ¦ Time Remaining, Current patient dialysis session ¦ Machine Configuration ¦ Dialysis | Duration | Time Remaining, Current patient dialysis session | Machine Configuration | Dialysis |  | HEMODIALYSIS | Dialysis Device State | Remaining Treatment Time | Number of minutes left for current dialysis session | Intradialytic | Status | Periodic | Numeric | XXXX | minutes | min | MDC_DIM_MIN | 1..1 | C10 |  | -- | -- | 0-9999 | 600 | 3 |  | MMM | 2 |  | 27652 | 158724
. . . MDC_TEMP_ROOM | Temperature ¦ ¦ Room ¦ Ambient | Temperature |  | Room | Ambient |  |  | Dialysis Device State | Room Temperature | Room Temperature in which the Dialysis Machine is operating | All | Status | Periodic | Numeric | XX.X | oC | Cel | MDC_DIM_DEGC | 0..1 | O |  | -- | -- | 0-99.9 | 22.1 | P |  | MMM | 2 | 57436 | 57436 | 188508
. . MDC_DEV_HDIALY_ANTICOAG_PUMP_CHAN | Subsystem ¦ ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Subsystem |  | Anticoagulant | Dialysis, Hemodialysis |  |  |  | Hemodialysis anticoagulant pump | The operational mode, status, profiles and settings related to the delivery of anticoagulants into the patient's blood. | [8 empty] | 1..* | C1 | [5 empty] | 3 |  | MVC (3) | 1 |  | 5407 | 70943
. . . MDC_ATTR_CHAN_NUM_LOGICAL | Number ¦ Pump number, sequence starting at 1, increments for each new pump ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Number | Pump number, sequence starting at 1, increments for each new pump | Anticoagulant | Dialysis, Hemodialysis |  | HEMODIALYSIS | Anticoagulant Pump | Anticoagulant Pump Number | The number associated with the anticoagulant pump in this channel. The value starts at 1 and is incremented for each anticoagulant pump in use. | Intradialytic | Parameter | Episodic | Numeric | X | N/A | -- | -- | 1..1 | M |  | -- | -- | 1-9 | 1 | 3 |  | 1 | 1 |  | 2606 | 68142
. . . MDC_HDIALY_ANTICOAG_MODE | Method (enumeration) ¦ Delivery Mode ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Method (enumeration) | Delivery Mode | Anticoagulant | Dialysis, Hemodialysis |  | HEMODIALYSIS | Anticoagulant Pump | Anticoagulation Mode | Method of anticoagulant administration during dialysis treatment | Intradialytic | Parameter | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 1..1 | C4 | _TBL_03 – Anticoagulation Mode | BOL CON BOLCON PRO BOLPRO NONE | Bolus – Administration of a fixed amount of anticoagulant drug during a specific time Continuous – Anticoagulant drug administered at a constant rate Combination of Bolus and Continuous mode Profile – Anticoagulant drug administered at a variable rate according to a programmed profile Combination of Bolus and Profile mode None | N/A | BOL | 3 |  | 1 | 2 |  | 27528 | 158600
. . . .Profile Parameters | [9 empty] | Contains the anticoagulant profile parameters. | [8 empty] | 0..1 | C20
. . . MDC_HDIALY_ANTICOAG_INFUS_RATE_SETTING | Flow Rate, Setting ¦ Infusion ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Flow Rate, Setting | Infusion | Anticoagulant | Dialysis, Hemodialysis |  | HEMODIALYSIS | Anticoagulant Pump | Anticoagulant Infusion Rate Setting | User defined amount of anticoagulation drug to be administered per unit time | Intradialytic | Parameter | Episodic | Numeric | XX.X | mL/h | mL/h | MDC_DIM_MILLI_L_PER_HR | 0..1 | C3 |  | -- | -- | 0-99.9 | 10.0 | 8
. . . MDC_HDIALY_ANTICOAG_INFUS_RATE | Flow Rate ¦ Infusion ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Flow Rate | Infusion | Anticoagulant | Dialysis, Hemodialysis |  | HEMODIALYSIS | Anticoagulant Pump | Anticoagulant Infusion Rate | User defined amount of anticoagulation drug to be administered per unit time | Intradialytic | Parameter | Episodic | Numeric | XX.X | mL/h | mL/h | MDC_DIM_MILLI_L_PER_HR | 0..1 | C3 |  | -- | -- | 0-99.9 | 10.0 | 3 |  | MMM | 2 |  | 27664 | 158736
. . . MDC_HDIALY_ANTICOAG_ACCUM_DELIV | Volume, Delivered ¦ Anticoagulant ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Volume, Delivered | Anticoagulant | Anticoagulant | Dialysis, Hemodialysis |  | HEMODIALYSIS | Anticoagulant Pump | Accumulated Delivered Anticoagulant | Total amount of anticoagulation drug administered to patient since the start of treatment | Intradialytic | Status | Both | Numeric | XXX.X | mL | mL | MDC_DIM_MILLI_L | 0..1 | C1 |  | -- | -- | 0-99.9 | 17.0 | 3 |  | MMM | 2 |  | 27656 | 158728
. . . MDC_HDIALY_ANTICOAG_BOLUS_VOL | Volume, Setting ¦ Anticoagulant, Bolus, Volume To Be Delivered ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Volume, Setting | Anticoagulant, Bolus, Volume To Be Delivered | Anticoagulant | Dialysis, Hemodialysis |  | HEMODIALYSIS | Anticoagulant Pump | Anticoagulant Bolus Volume | User defined amount of anticoagulation drug to be administered in this single large dose | Intradialytic | Parameter | Episodic | Numeric | XX.X | mL | mL | MDC_DIM_MILLI_L | 0..1 | C2 |  | -- | -- | 0-99.9 | 10.0 | 3 |  | MMM | 2 |  | 27660 | 158732
. . . MDC_HDIALY_INFUS_TIME_SETTING | Time ¦ Infusion ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Time | Infusion | Anticoagulant | Dialysis, Hemodialysis |  | HEMODIALYSIS | Anticoagulant Pump | Anticoagulant Infustion Time Setting | User specified infusion time. | Intradialytic | Parameter | Episodic | Numeric | XXXX | minutes | min | MDC_DIM_MIN | 0..1 | O |  | -- | -- | 0-9999 | 60 | 8
. . . MDC_EVT_HDIALY_ANTICOAG_STOP | StatusEvent ¦ Pump, Anticoagulant, ‹Stopped› ¦ FunctionalStatus ¦ Dialysis, Hemodialysis | StatusEvent | Pump, Anticoagulant, ‹Stopped› | FunctionalStatus | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Anticoagulant Pump | Anticoagulant Pump Stop | Notification that the anticoagulant pump has stopped | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1628 | 198236
. . . MDC_HDIALY_ANTICOAG_NAME | Identifier ¦ Anticoagulant Name ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Identifier | Anticoagulant Name | Anticoagulant | Dialysis, Hemodialysis |  | HEMODIALYSIS | Anticoagulant Pump | Anticoagulant Type | Name of the anticoagulant | Intradialytic | Parameter | Episodic | String | Alphanumeric | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | heparin | 3 |  | 1 | 2 |  | 27527 | 158599
. . . MDC_EVT_HDIALY_ANTICOAG_SYRINGE_EMPTY | StatusEvent ¦ Syringe, ‹Empty› ¦ Anticoagulant ¦ Dialysis, Hemodialysis | StatusEvent | Syringe, ‹Empty› | Anticoagulant | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Anticoagulant Pump | Syringe Empty | Notification that the syringe plunger has reached the end of its functional range | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1630 | 198238
. . . MDC_HDIALY_ANTICOAG_SYRINGE_BRAND | Identifier ¦ Syringe ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Identifier | Syringe | Anticoagulant | Dialysis, Hemodialysis |  | HEMODIALYSIS | Anticoagulant Pump | Syringe Name | Name of the syringe | Intradialytic | Parameter | Episodic | String | Alphanumeric | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | Fishman | 3 |  | 1 | 2 |  | 27530 | 158602
. . . MDC_HDIALY_ANTICOAG_SYRINGE_VOL | Volume ¦ Syringe, Initial Volume ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Volume | Syringe, Initial Volume | Anticoagulant | Dialysis, Hemodialysis |  | HEMODIALYSIS | Anticoagulant Pump | Syringe Volume | Volume of the syringe with anticoagulant. | Intradialytic | Parameter | Episodic | Numeric | XXX | mL | mL | MDC_DIM_MILLI_L | 0..1 | O |  | -- | -- | 0-999 | 60 | 3 |  | 1 | 2 |  | 27531 | 158603
. . . MDC_EVT_HDIALY_ANTICOAG_SYRINGE_SIZE | ErrorEvent ¦ Syringe, ‹Size Incorrect› ¦ Anticoagulant ¦ Dialysis, Hemodialysis | ErrorEvent | Syringe, ‹Size Incorrect› | Anticoagulant | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Anticoagulant Pump | Wrong Syringe Size | Notification that the size of the syringe is not the size defined for anticoagulant administration | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1632 | 198240
. . . MDC_HDIALY_ANTICOAG_LOCATION | Location (enumeration) ¦ Infusion ¦ Anticoagulant ¦ Dialysis, Hemodialysis | Location (enumeration) | Infusion | Anticoagulant | Dialysis, Hemodialysis |  | HEMODIALYSIS | Anticoagulant Pump | Anticoagulant Deliver Location | Location for delivery of the anticoagulant. | Intradialytic | Parameter | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 0..1 | O | _TBL_16 | PREP PREF POSTF PAT |  | N/A | PrePump | 8
. . MDC_DEV_HDIALY_BLOOD_PUMP_CHAN | Subsystem ¦ ¦ Blood Pump ¦ Dialysis, Hemodialysis | Subsystem |  | Blood Pump | Dialysis, Hemodialysis |  |  |  | Hemodialysis patient blood pump | The observations and settings related to the pressure, flow rate, temperature and conductivity of the patient's blood. | [8 empty] | 1..1 | M | [5 empty] | 3 |  | MVC (3) | 1 |  | 5411 | 70947
. . . MDC_HDIALY_BLD_PUMP_BLOOD_FLOW_RATE_SETTING | Flow Rate, Setting ¦ Blood ¦ Blood Pump ¦ Dialysis, Hemodialysis | Flow Rate, Setting | Blood | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Blood Flow Rate Setting | The rate at which the user programmed the blood flow | Intradialytic | Parameter | Episodic | Numeric | XXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 1..1 | M |  | -- | -- | 0-999 | 250 | 3 |  | MMM (0) | 258 |  | 27668 | 16935956
. . . MDC_HDIALY_BLD_PUMP_BLOOD_FLOW_RATE | Flow Rate, Current ¦ Blood ¦ Blood Pump ¦ Dialysis, Hemodialysis | Flow Rate, Current | Blood | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Actual Blood Flow Rate | Adjusted blood flow rate based on the blood flow rate setting and pressure drop caused by blood line, needle and vascular access | Intradialytic | Status | Periodic | numeric | XXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 0..1 | O |  | -- | -- | 0-999 | 250 | 3 |  | MMM | 2 |  | 27668 | 158740
. . . MDC_HDIALY_BLD_PUMP_BLOOD_FLOW_RATE_MEAN | Flow Rate, Mean ¦ Blood ¦ Blood Pump ¦ Dialysis, Hemodialysis | Flow Rate, Mean | Blood | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Average Blood Flow Rate | Average of the actual blood flow rate over the course of treatment. | Intradialytic | Status | Periodic | Numeric | XXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 0..1 | O |  | -- | -- | 0-999 | 250 | 3 |  | MMM (3) | 2 |  | 27671 | 158743
. . . MDC_HDIALY_BLD_PRESS_ART | Pressure ¦ Arterial, Pre-Pump ¦ Blood Pump ¦ Dialysis, Hemodialysis | Pressure | Arterial, Pre-Pump | Blood Pump | Dialysis, Hemodialysis | phys tech high low thr | HEMODIALYSIS | Blood Pump | Arterial Pressure | Pressure of arterial access line pre blood pump | Intradialytic | Status | Periodic | Numeric | ±XXX | mmHg | mm[Hg] | MDC_DIM_MMHG | 1..1 | M |  | -- | -- | -999 to 999 | -75 | 3 |  | MMM | 2 |  | 27672 | 158744
. . . MDC_HDIALY_BLD_PUMP_MODE | Pump Mode (enumeration) ¦ Blood ¦ Blood Pump ¦ Dialysis, Hemodialysis | Pump Mode (enumeration) | Blood | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Blood Pump Mode | Therapy method in which blood is retrieved and returned to the patient. | Intradialytic | Parameter | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 1..1 | M | _TBL_05 – Blood Pump Mode | 2N 1N1P 1N2P | Double Needles – Separate needles/catheter lumens for arterial and venous blood flow Single Needle/Single Pumps – The same needle/catheter lumen used for both arterial and venous blood flow, blood flow controlled by one pump Single Needle/Double Pumps – The same needle/catheter lumen used for both arterial and venous blood flow, blood flow controlled by two pumps | N/A | 2N | 3 |  | 1 | 2 |  | 27532 | 158604
. . . MDC_EVT_HDIALY_BLD_PUMP_STOP | StatusEvent ¦ Pump, ‹Stopped› ¦ Blood Pump ¦ Dialysis, Hemodialysis | StatusEvent | Pump, ‹Stopped› | Blood Pump | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Blood Pump | Blood Pump Stop | Notification that the blood pump has stopped | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | M |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1634 | 198242
. . . MDC_HDIALY_BLD_PUMP_TUBING_SIZE | Size ¦ Blood Pump Segment ¦ Blood Pump ¦ Dialysis, Hemodialysis | Size | Blood Pump Segment | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Blood Tubing Size | Inner diameter of pump segment of tubing used in extracorporeal circulation | Intradialytic | Parameter | Episodic | String | alphanumeric | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | 8 | 3 |  | 1 | 2 |  | 27533 | 158605
. . . MDC_HDIALY_BLOOD_TEMP_ART | Temperature ¦ Arterial, Pre-Pump ¦ Blood Pump ¦ Dialysis, Hemodialysis | Temperature | Arterial, Pre-Pump | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Arterial Blood Temperature | Arterial blood temperature as measured by the device. | Intradialytic | Status | Periodic | Numeric | XX.X | oC | Cel | MDC_DIM_DEGC | 0..1 | O |  | -- | -- | 0-99.9 | 39.1 | 3 |  | MMM | 2 |  | 27676 | 158748
. . . MDC_HDIALY_BLD_PUMP_CHANGE_IN_ENERGY | Energy ¦ Patient blood, to and from patient ¦ Blood Pump ¦ Dialysis, Hemodialysis | Energy | Patient blood, to and from patient | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Change In Energy | Measurement of the energy (temperate) flux to and from the patient | Intradialytic | Status | Periodic | Numeric | ±XXX.X | kJ/h | kJ/h | MDC_DIM_KILO_JOULES_PER_HR | 0..1 | O |  | -- | -- | -999.9 to 999.9 | 1.0 | 3 |  | MMM | 2 |  | 27680 | 158752
. . . MDC_HDIALY_BLD_PUMP_PRESS_ART_POST_PUMP | Pressure ¦ Arterial, Post-Pump ¦ Blood Pump ¦ Dialysis, Hemodialysis | Pressure | Arterial, Post-Pump | Blood Pump | Dialysis, Hemodialysis | tech high low thr | HEMODIALYSIS | Blood Pump | Post Pump Arterial Pressure | Pressure of arterial access line post blood pump | Intradialytic | Status | Periodic | Numeric | ±XXX | mmHg | mm[Hg] | MDC_DIM_MMHG | 0..1 | O |  | -- | -- | -999 to 999 | 100 | 3 |  | MMM | 2 |  | 27684 | 158756
. . . MDC_HDIALY_BLD_PUMP_PRIMING_VOL | Volume ¦ Extracorporeal Circuit, Priming ¦ Blood Pump ¦ Dialysis, Hemodialysis | Volume | Extracorporeal Circuit, Priming | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Priming Volume | Volume of fluid used to prime the extracorporeal circuit | Intradialytic | Parameter | Episodic | Numeric | XXX | mL | mL | MDC_DIM_MILLI_L | 0..1 | O |  | -- | -- | 0 to 999 | 191 | 3 |  | MMM | 2 |  | 27688 | 158760
. . . MDC_HDIALY_BLD_PUMP_SINGLE_NEEDLE_PRESS | Pressure ¦ Threshold, used to switch phases at high pressure ¦ Blood Pump ¦ Dialysis, Hemodialysis | Pressure | Threshold, used to switch phases at high pressure | Blood Pump | Dialysis, Hemodialysis | tech high low thr | HEMODIALYSIS | Blood Pump | Single Needle Switching Pressure | In single needle mode, the pressure used to switch phases at high pressure | Intradialytic | Parameter | Episodic | Numeric | ± XXX | mmHg | mm[Hg] | MDC_DIM_MMHG | 0..1 | O |  | -- | -- | 0 to 999 | 200 | 3 |  | MMM | 2 |  | 27692 | 158764
. . . MDC_HDIALY_BLD_PUMP_STROKE_VOL | Volume ¦ Single Needle Stroke Cycle ¦ Blood Pump ¦ Dialysis, Hemodialysis | Volume | Single Needle Stroke Cycle | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Stroke Volume | The volume of blood passing through the dialyzer during each single needle cycle. | Intradialytic | Status | Periodic | Numeric | XXX | mL | mL | MDC_DIM_MILLI_L | 0..1 | C5 |  | -- | -- | 0 to 999 | 10 | 3 |  | MMM | 2 |  | 27696 | 158768
. . . MDC_HDIALY_BLD_PUMP_BLOOD_PROCESSED_TOTAL | Volume, Processed ¦ Blood ¦ Blood Pump ¦ Dialysis, Hemodialysis | Volume, Processed | Blood | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Total Blood Processed | Accumulated volume of blood circulated by extracorporeal circulation. | Intradialytic | Status | Periodic | Numeric | XXXX.XX | L | L | MDC_DIM_L | 0..1 | O |  | -- | -- | 0 to 9999.99 | 120 | 3 |  | MMM | 2 |  | 27700 | 158772
. . . MDC_HDIALY_BLD_PUMP_PRESS_VEN | Pressure ¦ Venous ¦ Blood Pump ¦ Dialysis, Hemodialysis | Pressure | Venous | Blood Pump | Dialysis, Hemodialysis | phys tech high low thr | HEMODIALYSIS | Blood Pump | Venous Pressure | Pressure of the venous access line | Intradialytic | Status | Periodic | Numeric | ±XXX | mmHg | mm[Hg] | MDC_DIM_MMHG | 1..1 | M |  | -- | -- | -999 to 999 | 200 | 3 |  | MMM | 2 |  | 27704 | 158776
. . . MDC_HDIALY_BLOOD_TEMP_VEN | Temperature ¦ Venous ¦ Blood Pump ¦ Dialysis, Hemodialysis | Temperature | Venous | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Venous Temperature | Temperature of the blood measured in the venous access line | Intradialytic | Status | Periodic | Numeric | XX.X | oC | Cel | MDC_DIM_DEGC | 0..1 | O |  | -- | -- | 0 to 99.9 | 39.1 | 3 |  | MMM | 2 |  | 27708 | 158780
. . . MDC_HDIALY_BLOOD_TEMP_VEN_SETTING | Temperature, Setting ¦ Venous ¦ Blood Pump ¦ Dialysis, Hemodialysis | Temperature, Setting | Venous | Blood Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Blood Pump | Venous Temperature Setting | Desired temperature of the blood in the venous access line | Intradialytic | Status | Periodic | Numeric | XX.X | oC | Cel | MDC_DIM_DEGC | 0..1 | O |  | -- | -- | 0 to 99.9 | 39.0 | 3 |  | MMM | 258 |  | 27708 | 16935996
. . MDC_DEV_HDIALY_FLUID_CHAN | Subsystem ¦ ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Subsystem |  | Fluid Pump | Dialysis, Hemodialysis |  |  |  | Dialysate | The observations and settings related to fluid chemistry and pressures, flow rates, temperature and conductivity of the dialysate solution. | [8 empty] | 1..1 | C6 | [5 empty] | 3 |  | MVC (3) | 1 |  | 5415 | 70951
. . . MDC_HDIALY_DIALYSATE_FLOW_MODE | Pump Mode (enumeration) ¦ Dialysate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Pump Mode (enumeration) | Dialysate | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Flow Mode | The mode of controlling dialysate flow rate | All | Parameter | Episodic | Enumeration | alphanumeric | N/A | -- | -- | 1..1 | C6 | _TBL_06 – Dialysis Flow Mode | CONST AUTO PRO STBY NONE | Constant dialysis fluid flow Auto – Dialysis fluid flow is automatically controlled by the dialysis machine Profile – Variable dialysis fluid flow according to a programmed profile Standby – dialysis fluid flow is reduced when dialysis machine is not providing treatment Dialysis fluid does not flow. | N/A | CONST | 3 |  | 1 | 2 |  | 27534 | 158606
. . . .Profile Parameters | [9 empty] | Contains the dialysate profile parameters. | [8 empty] | 0..1 | C22
. . . MDC_HDIALY_DIALYSATE_FLOW_RATE_SETTING | Flow Rate, Setting ¦ Dialysate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Flow Rate, Setting | Dialysate | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Flow Rate Setting | Rate at which the user programmed the dialysate flow | Intradialytic | Parameter | Episodic | Numeric | XXXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 0..1 | C6 |  | -- | -- | 0 to 9999 | 100 | 3 |  | MMM | 258 |  | 27720 | 16936008
. . . MDC_HDIALY_DIALYSATE_FLOW_RATE | Flow Rate, Current ¦ Dialysate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Flow Rate, Current | Dialysate | Fluid Pump | Dialysis, Hemodialysis | tech low | HEMODIALYSIS | Dialysis Fluid | Actual Dialysate Flow Rate | Rate of dialysate flow at any given point during a treatment | Intradialytic | Status | Periodic | Numeric | XXXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 1..1 | C6 |  | -- | -- | 0 to 9999 | 99 | 3 |  | MMM | 2 |  | 27720 | 158792
. . . MDC_HDIALY_DIALYSATE_VOL_DELIV | Volume, Delivered ¦ Dialysate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Volume, Delivered | Dialysate | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Total Dialysate Volume | Total amount of dialysate used during a treatment | Intradialytic | Status | Periodic | Numeric | XXX.XX | L | L | MDC_DIM_L | 0..1 | O |  | -- | -- | 0 to 999.99 | 24.34 | 3 |  | MMM | 2 |  | 27776 | 158848
. . . MDC_HDIALY_DIALYSATE_VOL_DELIV_SETTING | Volume, Setting ¦ Dialysate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Volume, Setting | Dialysate | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Volume Setting | User programmed dialysate volume to deliver | Intradialytic | Parameter | Episodic | Numeric | XXX.XX | L | L | MDC_DIM_L | 0..1 | O |  | -- | -- | 0 to 999.99 | 60 | 8
. . . MDC_HDIALY_DIALYSATE_NAME | Identifier ¦ Dialysate Name ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Identifier | Dialysate Name | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Name | String containing a descriptive name of the dialysate | Intradialytic | Parameter | Episodic | String | Alphanumeric | N/A | -- | -- | 1..1 | O |  | -- | -- | N/A | FMC smartbag 111.5 | 3 |  | 1 | 2 |  | 27536 | 158608
. . . MDC_HDIALY_BICARB_COND | Conductivity ¦ Bicarbonate (HCO3) ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Conductivity | Bicarbonate (HCO3) | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Actual Bicarbonate Conductivity | Measured conductivity of bicarbonate | Intradialytic | Status | Periodic | Numeric | XX.XX | mS/cm | mS/cm | MDC_DIM_MILLI_SIEMENS_PER_CM | 0..1 | O |  | -- | -- | 0 to 99.99 | 13.81 | 3 |  | MMM | 2 |  | 27712 | 158784
. . . MDC_HDIALY_BICARB_COND_SETTING | Conductivity, Setting ¦ Bicarbonate (HCO3) ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Conductivity, Setting | Bicarbonate (HCO3) | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Bicarbonate Conductivity Setting | Bicarbonate conductivity | Intradialytic | Parameter | Episodic | Numeric | XX.XX | mS/cm | mS/cm | MDC_DIM_MILLI_SIEMENS_PER_CM | 0..1 | O |  | -- | -- | 0 to 99.99 | 13.81 | 3 |  | MMM | 258 |  | 27712 | 16936000
. . . MDC_HDIALY_DIALYSATE_COND | Conductivity ¦ Dialysate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Conductivity | Dialysate | Fluid Pump | Dialysis, Hemodialysis | tech high low | HEMODIALYSIS | Dialysis Fluid | Actual Conductivity | Measured conductivity of dialysate | Intradialytic | Status | Periodic | Numeric | XX.XX | mS/cm | mS/cm | MDC_DIM_MILLI_SIEMENS_PER_CM | 1..1 | C12 |  | -- | -- | 0 to 99.99 | 13.81 | 3 |  | MMM | 2 |  | 27716 | 158788
. . . MDC_HDIALY_DIALYSATE_COND_SETTING | Conductivity, Setting ¦ Dialysate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Conductivity, Setting | Dialysate | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Conductivity Setting | Acid and Bicarbonate conductivity combined | Intradialytic | Parameter | Episodic | Numeric | XX.XX | mS/cm | mS/cm | MDC_DIM_MILLI_SIEMENS_PER_CM | 0..1 | O |  | -- | -- | 0 to 99.99 | 13.81 | 3 |  | MMM | 258 |  | 27716 | 16936004
. . . MDC_HDIALY_DIALYSATE_TEMP | Temperature, Current ¦ Dialysate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Temperature, Current | Dialysate | Fluid Pump | Dialysis, Hemodialysis | tech high low | HEMODIALYSIS | Dialysis Fluid | Actual Dialysate Temperature | Temperature of dialysate | Intradialytic | Status | Periodic | Numeric | XX.X | °C | Cel | MDC_DIM_DEGC | 1..1 | O |  | -- | -- | 0 to 99.9 | 41.0 | 3 |  | MMM | 2 |  | 27724 | 158796
. . . MDC_HDIALY_DIALYSATE_TEMP_SETTING | Temperature, Setting ¦ Dialysate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Temperature, Setting | Dialysate | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Temperature Setting | Temperature at which dialysate is to be delivered to patient | Intradialytic | Parameter | Episodic | Numeric | XX.X | °C | Cel | MDC_DIM_DEGC | 1..1 | O |  | -- | -- | 0 to 99.9 | 41.1 | 3 |  | MMM | 258 |  | 27724 | 16936012
. . . MDC_HDIALY_DIALYSATE_AMMONIA | Concentration ¦ Ammonia ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration | Ammonia | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Ammonia Concentration | Measured ammonia concentration of the dialysate | Intradialytic | Status | Periodic | Numeric | X.XX | ppm | [ppm] | MDC_DIM_PARTS_PER_10_TO_6 | 0..1 | O |  | -- | -- | 0 to 9.99 | 0.00 | 3 |  | MMM | 2 |  | 27728 | 158800
. . . MDC_HDIALY_DIALYSATE_FLOW_RATE_MEAN | Flow Rate, Mean ¦ Dialysate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Flow Rate, Mean | Dialysate | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Average Dialysate Flow Rate | Average rate of dialysate flow during course of treatment | Intradialytic | Status | Periodic | Numeric | XXXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 0..1 | O |  | -- | -- | 0 to 9999 | 100 | 3 |  | MMM (3) | 2 |  | 27723 | 158795
. . . MDC_HDIALY_CONC_HCO3_SETTING | Concentration, Setting ¦ Bicarbonate (HCO3) ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration, Setting | Bicarbonate (HCO3) | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Bicarbonate Concentration Setting | User programmed Bicarbonate concentration in dialysate. | Intradialytic | Parameter | Episodic | Numeric | XXX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 1..1 | O |  | -- | -- | 0 to 999.9 | 0 | 3 |  | MMM | 258 |  | 27732 | 16936020
. . . MDC_HDIALY_CONC_HCO3_MODE | Delivery Mode (enumeration) ¦ Bicarbonate (HCO3) ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Delivery Mode (enumeration) | Bicarbonate (HCO3) | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Bicarbonate Mode | Process for delivering adjusted bicarbonate concentration in dialysate | Intradialytic | Parameter | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 1..1 | O | _TBL_14 – Bicarb Delivery Mode | CONST PRO NA | Constant – Bicarbonate concentration in dialysis fluid is constant. Profile – Bicarbonate concentration in dialysis fluid is adjusted according to a programmed profile Not Applicable | N/A | PRO | 3 |  | 1 | 2 |  | 27535 | 158607
. . . MDC_EVT_HDIALY_BLOOD_LEAK | LimitEvent ¦ Blood In Dialysate, ‹Detected› ¦ Fluid Pump ¦ Dialysis, Hemodialysis | LimitEvent | Blood In Dialysate, ‹Detected› | Fluid Pump | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Dialysis Fluid | Blood Leak | Notification that blood has been detected in the dialysate fluid | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | M |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1636 | 198244
. . . MDC_HDIALY_DIALYSATE_CONC_ACETATE | Concentration ¦ Acetate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration | Acetate | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Acetate Concentration | Acetate concentration in dialysate | Intradialytic | Parameter | Episodic | Numeric | XX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 99.9 | 3.00 | 3 |  | MMM | 2 |  | 27736 | 158808
. . . MDC_HDIALY_DIALYSATE_CONC_HCO3 | Concentration ¦ Bicarbonate (HCO3) ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration | Bicarbonate (HCO3) | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Bicarbonate Concentration | Bicarbonate concentration in dialysate | Intradialytic | Parameter | Episodic | Numeric | XXX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 999.9 | 0.0 | 3 |  | MMM | 2 |  | 27740 | 158812
. . . MDC_HDIALY_DIALYSATE_CONC_CHLORIDE | Concentration ¦ Chloride ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration | Chloride | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Chloride Concentration | Chloride concentration in dialysate | Intradialytic | Parameter | Episodic | Numeric | XXX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 999 | 108.0 | 3 |  | MMM | 2 |  | 27744 | 158816
. . . MDC_HDIALY_DIALYSATE_CONC_MG | Concentration ¦ Acetate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration | Acetate | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Magnesium Concentration | Magnesium concentration in dialysate | Intradialytic | Parameter | Episodic | Numeric | XX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 99.9 | 0.5 | 3 |  | MMM | 2 |  | 27748 | 158820
. . . MDC_HDIALY_DIALYSATE_CONC_CA | Concentration ¦ Calcium (CA) ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration | Calcium (CA) | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Calcium Concentration | Calcium concentration in dialysate | Intradialytic | Parameter | Episodic | Numeric | X.XX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 9.99 | 1.50 | 3 |  | MMM | 2 |  | 27752 | 158824
. . . MDC_HDIALY_DIALYSATE_CONC_CITRATE | Concentration ¦ Citrate ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration | Citrate | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Citrate Concentration | Citrate concentration in dialysate | Intradialytic | Parameter | Episodic | Numeric | XX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 99.9 | 0.0 | 3 |  | MMM | 2 |  | 27756 | 158828
. . . MDC_HDIALY_DIALYSATE_CONC_GLU | Concentration ¦ Glucose ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration | Glucose | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Glucose Concentration | Glucose concentration in dialysate | Intradialytic | Parameter | Episodic | Numeric | XXX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 999 | 1.00 | 3 |  | MMM | 2 |  | 27760 | 158832
. . . MDC_HDIALY_DIALYSATE_CONC_K | Concentration ¦ Potassium ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration | Potassium | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Potassium Concentration | Potassium concentration in dialysate | Intradialytic | Parameter | Episodic | Numeric | X.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 9.9 | 1.00 | 3 |  | MMM | 2 |  | 27764 | 158836
. . . MDC_HDIALY_DIALYSATE_CONC_NA | Concentration ¦ Sodium (NA) ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration | Sodium (NA) | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Sodium Concentration | Sodium concentration in dialysate | Intradialytic | Parameter | Episodic | Numeric | XXXX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 9999 | 138 | 3 |  | MMM | 2 |  | 27768 | 158840
. . . MDC_HDIALY_DIALYSATE_CONC_NA_SETTING | Concentration, Setting ¦ Sodium (NA) ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration, Setting | Sodium (NA) | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Sodium concentration Setting | User programmed sodium concentration in dialysate | Intradialytic | Parameter | Episodic | Numeric | XXXX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 1..1 | O |  | -- | -- | 0 to 9999 | 140 | 3 |  | MMM | 258 |  | 27768 | 16936056
. . . MDC_HDIALY_DIALYSATE_CONC_NA_MODE | Pump Mode (enumeration) ¦ Sodium (NA) ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Pump Mode (enumeration) | Sodium (NA) | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | Dialysate Sodium Mode | Process for delivering adjusted sodium concentration in dialysate | Intradialytic | Parameter | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 1..1 | O | _TBL_15 – Sodium Delivery Mode | CONST PRO | Constant – Sodium concentration in dialysis fluid is constant Profile – Sodium concentration in dialysis fluid is adjusted according to a programmed profile | N/A | CONST | 3 |  | 1 | 2 |  | 27537 | 158609
. . . .Profile Parameters | [9 empty] | Contains the sodium delivery profile parameters. | [8 empty] | 0..1 | C23
. . . MDC_HDIALY_DIALYSATE_CONC_PH | Concentration ¦ pH (H+ logarithmic) ¦ Fluid Pump ¦ Dialysis, Hemodialysis | Concentration | pH (H+ logarithmic) | Fluid Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Dialysis Fluid | pH | Measured pH of the dialysate | Intradialytic | Status | Periodic | Numeric | XX.X | pH | [pH] | MDC_DIM_PH | 0..1 | O |  | -- | -- | 0 to 9.99 | 7.0 | 3 |  | MMM | 2 |  | 27772 | 158844
. . MDC_DEV_HDIALY_FILTER_CHAN | Subsystem ¦ ¦ Filter ¦ Dialysis, Hemodialysis | Subsystem |  | Filter | Dialysis, Hemodialysis |  |  |  | Hemodialysis filter | The identity and observations related to the hemodialysis filter, the replacable semi-permeable membrane filter that removes toxins from the blood and transport of other solutes by diffusion and/or convection. | [8 empty] | 1..1 | M | [5 empty] | 3 |  | MVC (3) | 1 |  | 5419 | 70955
. . . MDC_HDIALY_FILTER_NAME | Identifier ¦ Dialyzer Filter Name (manufacturer and model) ¦ Filter ¦ Dialysis, Hemodialysis | Identifier | Dialyzer Filter Name (manufacturer and model) | Filter | Dialysis, Hemodialysis |  | HEMODIALYSIS | Filter (Dialyzer) | Dialyzer Name | Dialyzer manufacturer and model | Intradialytic | Parameter | Episodic | String | Alphanumeric | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | "NxStage CAR125" | 3 |  | 1 | 2 |  | 27538 | 158610
. . . MDC_HDIALY_FILTER_TRANSMEMBRANE_PRESS | Pressure ¦ Transmembrane (Difference between blood compartment and dialysate compartment) ¦ Filter ¦ Dialysis, Hemodialysis | Pressure | Transmembrane (Difference between blood compartment and dialysate compartment) | Filter | Dialysis, Hemodialysis | tech high low | HEMODIALYSIS | Filter (Dialyzer) | Actual Transmembrane Pressure | Difference in pressure between blood compartment and dialysate compartment of dialyzer | Intradialytic | Status | Periodic | Numeric | ±XXX | mmHg | mm[Hg] | MDC_DIM_MMHG | 1..1 | M |  | -- | -- | -999 to 999 | 35 | 3 |  | MMM | 2 |  | 27780 | 158852
. . . MDC_HDIALY_FILTER_NUM | Number ¦ Filter number, sequence starting at 1, increments for each new filter ¦ Filter ¦ Dialysis, Hemodialysis | Number | Filter number, sequence starting at 1, increments for each new filter | Filter | Dialysis, Hemodialysis |  | HEMODIALYSIS | Filter (Dialyzer) | Filter Number | The filter number since the start of treatment. This value starts at 1 and increments with each new filter used during the current treatment. | Intradialytic | Status | Periodic | Numeric | X | N/A | -- | -- | 0..1 | O |  | -- | -- | 1 to 9 | 1 | 3 |  | 1 | 2 |  | 27539 | 158611
. . . MDC_HDIALY_FILTER_UDI | Identifier ¦ Current Filter UDI ¦ Filter ¦ Dialysis, Hemodialysis | Identifier | Current Filter UDI | Filter | Dialysis, Hemodialysis |  | HEMODIALYSIS | Filter (Dialyzer) | Filter UDI | The UDI for the current filter. | Intradialytic | Status | Periodic | String | N/A | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | 01008422891018381720120010LOT00604 | 3 |  | 1 | 2 |  | 27540 | 158612
. . MDC_DEV_HDIALY_CONVECTIVE_CHAN | Subsystem ¦ ¦ Convective Pump ¦ Dialysis, Hemodialysis | Subsystem |  | Convective Pump | Dialysis, Hemodialysis |  |  |  | Hemodialysis convective channel | The observations and settings related to convective clearance of solutes out of the blood compartment along with movement of water and/or ultrafiltration. | [8 empty] | 1..1 | C8 | [5 empty] | 3 |  | MVC (3) | 1 |  | 5423 | 70959
. . . MDC_HDIALY_RF_DILUTION_LOCATION | Location (enumeration) ¦ Replacement Fluid, Dilution Location ¦ Convective Pump ¦ Dialysis, Hemodialysis | Location (enumeration) | Replacement Fluid, Dilution Location | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Replacement Fluid Dilution Location | Delivery location of the replacement fluid | Intradialytic | Parameter | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 1..1 | C8 | _TBL_08 – RF Dilution Location | PREF POSTF PREF-POSTF PREP-POSTF | Pre-Filter Post-Filter Both Pre-Filter and Post-Filter Both Pre-Pump and Post-Filter | N/A | PREF | 3 |  | 1 | 2 |  | 27541 | 158613
. . . MDC_HDIALY_RF_FLOW_MODE | Pump Mode (enumeration) ¦ Replacement Fluid ¦ Convective Pump ¦ Dialysis, Hemodialysis | Pump Mode (enumeration) | Replacement Fluid | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Replacement Fluid Flow Mode | The mode of controlling replacement fluid flow rate | All | Parameter | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 1..1 | C8 | _TBL_07 – Replacement Fluid Delivery Mode | CONST PRO NONE BOLUS AUTO | Constant replacement fluid flow Profile – Variable replacement fluid flow according to a programmed profile None, Not Applicable. Bolus – administration of a fixed amount of replacement fluid during a specific time Automatic – Replacement fluid flow is automatically controlled by the dialysis machine | N/A | CONST | 3 |  | 1 | 2 |  | 27544 | 158616
. . . .Profile Parameters | [9 empty] | Contains the replacement fluid delivery profile parameters. | [8 empty] | 0..1 | C24
. . . MDC_HDIALY_RF_POST_FILTER_FLUID_NAME | Identifier ¦ Replacement Fluid, Post Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Identifier | Replacement Fluid, Post Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter RF Solution Name | String containing a descriptive name of the post-filter replacement fluid | Intradialytic | Parameter | Episodic | String | Alphanumeric | N/A | -- | -- | 1..1 | O |  | -- | -- | N/A | multiBIC | 3 |  | 1 | 2 |  | 27542 | 158614
. . . MDC_HDIALY_RF_POST_FILTER_CONC_ACETATE | Concentration ¦ Acetate, Replacement Fluid, Post-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Acetate, Replacement Fluid, Post-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter RF Acetate concentration | Acetate concentration in the post-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 99.9 | 0 | 3 |  | MMM | 2 |  | 27784 | 158856
. . . MDC_HDIALY_RF_POST_FILTER_CONC_HCO3 | Concentration ¦ Bicarbonate (HCO3), Replacement Fluid, Post-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Bicarbonate (HCO3), Replacement Fluid, Post-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter RF Bicarbonate concentration | Bicarbonate concentration in the post-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XXX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 999.9 | 35 | 3 |  | MMM | 2 |  | 27788 | 158860
. . . MDC_HDIALY_RF_POST_FILTER_CONC_CHLORIDE | Concentration ¦ Chloride, Replacement Fluid, Post-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Chloride, Replacement Fluid, Post-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter RF Chloride concentration | Chloride concentration in the post-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XXX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 999 | 111 | 3 |  | MMM | 2 |  | 27792 | 158864
. . . MDC_HDIALY_RF_POST_FILTER_CONC_MG | Concentration ¦ Magnesium, Replacement Fluid, Post-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Magnesium, Replacement Fluid, Post-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter RF Magnesium concentration | Magnesium concentration in the post-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 99.9 | 0.5 | 3 |  | MMM | 2 |  | 27796 | 158868
. . . MDC_HDIALY_RF_POST_FILTER_CONC_CA | Concentration ¦ Calcium (CA), Replacement Fluid, Post-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Calcium (CA), Replacement Fluid, Post-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter RF Calcium concentration | Calcium concentration in the post-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | X.XX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 9.99 | 1.5 | 3 |  | MMM | 2 |  | 27800 | 158872
. . . MDC_HDIALY_RF_POST_FILTER_CONC_CITRATE | Concentration ¦ Citrate, Replacement Fluid, Post-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Citrate, Replacement Fluid, Post-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter RF Citrate concentration | Citrate concentration in the post-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 99.9 | 0 | 3 |  | MMM | 2 |  | 27804 | 158876
. . . MDC_HDIALY_RF_POST_FILTER_CONC_GLU | Concentration ¦ Glucose, Replacement Fluid, Post-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Glucose, Replacement Fluid, Post-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter RF Glucose concentration | Glucose concentration in the post-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XXX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 999 | 5.55 | 3 |  | MMM | 2 |  | 27808 | 158880
. . . MDC_HDIALY_RF_POST_FILTER_CONC_K | Concentration ¦ Potassium, Replacement Fluid, Post-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Potassium, Replacement Fluid, Post-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter RF Potassium concentration | Potassium concentration in the post-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | X.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 9.9 | 2 | 3 |  | MMM | 2 |  | 27812 | 158884
. . . MDC_HDIALY_RF_POST_FILTER_CONC_NA | Concentration ¦ Sodium (NA), Replacement Fluid, Post-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Sodium (NA), Replacement Fluid, Post-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter RF Sodium concentration | Sodium concentration in the post-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XXXX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 9999 | 140 | 3 |  | MMM | 2 |  | 27816 | 158888
. . . MDC_HDIALY_RF_POST_FILTER_FLOW_RATE_SETTING | Flow Rate, Setting ¦ Replacement Fluid, Pre-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Flow Rate, Setting | Replacement Fluid, Pre-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter Replacement Fluid Flow Rate Setting | Rate at which the user programmed the post -filter replacement fluid flow | Intradialytic | Parameter | Episodic | Numeric | XXXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 0..1 | C7 |  | -- | -- | 0 to 9999 | 60 | 3 |  | MMM | 258 |  | 27820 | 16936108
. . . MDC_HDIALY_RF_POST_FILTER_FLOW_RATE | Flow Rate, Current ¦ Replacement Fluid, Pre-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Flow Rate, Current | Replacement Fluid, Pre-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Actual Post-Filter Replacement Fluid Flow Rate | Rate of post -filter replacement fluid flow at any given point during a treatment | Intradialytic | Status | Periodic | Numeric | XXXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 1..1 | C7 |  | -- | -- | 0 to 9999 | 60 | 3 |  | MMM | 2 |  | 27820 | 158892
. . . MDC_HDIALY_RF_POST_FILTER_FLOW_RATE_MEAN | Flow Rate, Mean ¦ Replacement Fluid, Pre-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Flow Rate, Mean | Replacement Fluid, Pre-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Average Post-Filter Replacement Fluid Flow Rate | Average rate of post -filter replacement fluid flow during course of treatment. | Intradialytic | Status | Periodic | Numeric | XXXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 0..1 | O |  | -- | -- | 0 to 9999 | 60 | 3 |  | MMM (3) | 2 |  | 27823 | 158895
. . . MDC_HDIALY_RF_POST_FILTER_VOL_SETTING | Volume, Setting ¦ Replacement Fluid, Pre-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Volume, Setting | Replacement Fluid, Pre-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter Replacement Fluid Volume setting | Programmed total amount of post -filter replacement fluid volume | Intradialytic | Parameter | Episodic | Numeric | XXX.XX | L | L | MDC_DIM_L | 0..1 | O |  | -- | -- | 0 to 999.99 | 12.00 | 3 |  | MMM | 258 |  | 27828 | 16936116
. . . MDC_HDIALY_RF_POST_FILTER_VOL | Volume, Delivered ¦ Replacement Fluid, Pre-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Volume, Delivered | Replacement Fluid, Pre-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Total Post-Filter Replacement Fluid Volume | Total amount of post -filter replacement fluid used during a treatment | Intradialytic | Status | Periodic | Numeric | XXX.XX | L | L | MDC_DIM_L | 0..1 | O |  | -- | -- | 0 to 999.99 | 6.00 | 3 |  | MMM | 2 |  | 27828 | 158900
. . . MDC_HDIALY_RF_POST_FILTER_TEMP_SETTING | Temperature, Setting ¦ Replacement Fluid, Pre-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Temperature, Setting | Replacement Fluid, Pre-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Post-Filter Replacement Fluid Temperature Setting | Temperature at which post-filter replacement fluid is to be delivered to patient | Intradialytic | Parameter | Episodic | Numeric | XX.X | °C | Cel | MDC_DIM_DEGC | 1..1 | C7 |  | -- | -- | 0 to 99.9 | 41.0 | 3 |  | MMM | 258 |  | 27824 | 16936112
. . . MDC_HDIALY_RF_POST_FILTER_TEMP | Temperature ¦ Replacement Fluid, Pre-Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Temperature | Replacement Fluid, Pre-Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Actual Post-Filter Replacement Fluid Temperature | Temperature of post-filter replacement fluid | Intradialytic | Status | Periodic | Numeric | XX.X | °C | Cel | MDC_DIM_DEGC | 1..1 | C7 |  | -- | -- | 0 to 99.9 | 39.9 | 3 |  | MMM | 2 |  | 27824 | 158896
. . . MDC_HDIALY_RF_PRE_FILTER_FLUID_NAME | Identifier ¦ Replacement Fluid, Pre Pump ¦ Convective Pump ¦ Dialysis, Hemodialysis | Identifier | Replacement Fluid, Pre Pump | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter RF Solution Name | String containing a descriptive name of the pre-filter replacement fluid | Intradialytic | Parameter | Episodic | String | Alphanumeric | N/A | -- | -- | 1..1 | O |  | -- | -- | N/A | multiBIC | 3 |  | 1 | 2 |  | 27543 | 158615
. . . MDC_HDIALY_RF_PRE_FILTER_CONC_ACETATE | Concentration ¦ Acetate, Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Acetate, Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter RF Acetate Concentration | Acetate concentration in the pre-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 99.9 | 0 | 3 |  | MMM | 2 |  | 27832 | 158904
. . . MDC_HDIALY_RF_PRE_FILTER_CONC_HCO3 | Concentration ¦ Bicarbonate (HCO3), Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Bicarbonate (HCO3), Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter RF Bicarbonate Concentration | Bicarbonate concentration in the pre-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XXX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 999.9 | 35 | 3 |  | MMM | 2 |  | 27836 | 158908
. . . MDC_HDIALY_RF_PRE_FILTER_CONC_CHLORIDE | Concentration ¦ Chloride, Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Chloride, Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter RF Chloride Concentration | Chloride concentration in the pre-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XXX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 999 | 111 | 3 |  | MMM | 2 |  | 27840 | 158912
. . . MDC_HDIALY_RF_PRE_FILTER_CONC_MG | Concentration ¦ Magnesium, Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Magnesium, Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter RF Magnesium Concentration | Magnesium concentration in the pre-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 99.9 | 0.5 | 3 |  | MMM | 2 |  | 27844 | 158916
. . . MDC_HDIALY_RF_PRE_FILTER_CONC_CA | Concentration ¦ Calcium (CA), Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Calcium (CA), Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter RF Calcium Concentration | Calcium concentration in the pre-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | X.XX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 9.99 | 1.5 | 3 |  | MMM | 2 |  | 27848 | 158920
. . . MDC_HDIALY_RF_PRE_FILTER_CONC_CITRATE | Concentration ¦ Citrate, Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Citrate, Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter RF Citrate Concentration | Citrate concentration in the pre-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XX.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 99.9 | 0 | 3 |  | MMM | 2 |  | 27852 | 158924
. . . MDC_HDIALY_RF_PRE_FILTER_CONC_GLU | Concentration ¦ Glucose, Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Glucose, Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter RF Glucose Concentration | Glucose concentration in the pre-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XXX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 999 | 5.55 | 3 |  | MMM | 2 |  | 27856 | 158928
. . . MDC_HDIALY_RF_PRE_FILTER_CONC_K | Concentration ¦ Potassium (K), Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Potassium (K), Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter RF Potassium Concentration | Potassium concentration in the pre-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | X.X | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 9.9 | 2 | 3 |  | MMM | 2 |  | 27860 | 158932
. . . MDC_HDIALY_RF_PRE_FILTER_CONC_NA | Concentration ¦ Sodium (NA), Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Concentration | Sodium (NA), Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter RF Sodium Concentration | Sodium concentration in the pre-filter replacement fluid | Intradialytic | Parameter | Episodic | Numeric | XXXX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 9999 | 140 | 3 |  | MMM | 2 |  | 27864 | 158936
. . . MDC_HDIALY_RF_PRE_FILTER_FLOW_RATE_SETTING | Flow Rate, Setting ¦ Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Flow Rate, Setting | Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter Replacement Fluid Flow Rate Setting | Rate at which the user programmed the pre-filter replacement fluid flow | Intradialytic | Parameter | Episodic | Numeric | XXXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 0..1 | C8 |  | -- | -- | 0 to 9999 | 60 | 3 |  | MMM | 258 |  | 27868 | 16936156
. . . MDC_HDIALY_RF_PRE_FILTER_FLOW_RATE | Flow rate, Current ¦ Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Flow rate, Current | Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Actual Pre-Filter Replacement Fluid Flow Rate | Rate of pre-filter replacement fluid flow at any given point during a treatment | Intradialytic | Status | Periodic | Numeric | XXXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 1..1 | C7 |  | -- | -- | 0 to 9999 | 60 | 3 |  | MMM | 2 |  | 27868 | 158940
. . . MDC_HDIALY_RF_PRE_FILTER_FLOW_RATE_MEAN | Flow Rate, Mean ¦ Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Flow Rate, Mean | Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Average Pre-Filter Replacement Fluid Flow Rate | Average rate of pre-filter replacement fluid flow during course of treatment. | Intradialytic | Status | Periodic | Numeric | XXXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 0..1 | O |  | -- | -- | 0 to 9999 | 60 | 3 |  | MMM (3) | 2 |  | 27871 | 158943
. . . MDC_HDIALY_RF_PRE_FILTER_VOL_SETTING | Volume, Setting ¦ Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Volume, Setting | Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter Replacement Fluid Volume Setting | Programmed total amount of pre -filter replacement fluid volume | Intradialytic | Parameter | Episodic | Numeric | XXX.XX | L | L | MDC_DIM_L | 0..1 | O |  | -- | -- | 0 to 999.99 | 3.00 | 3 |  | MMM | 258 |  | 27876 | 16936164
. . . MDC_HDIALY_RF_PRE_FILTER_VOL | Volume, Delivered ¦ Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Volume, Delivered | Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Total Pre-Filter Replacement Fluid Volume | Total amount of pre-filter replacement fluid used during a treatment | Intradialytic | Status | Periodic | Numeric | XXX.XX | L | L | MDC_DIM_L | 0..1 | O |  | -- | -- | 0 to 999.99 | 1.23 | 3 |  | MMM | 2 |  | 27876 | 158948
. . . MDC_HDIALY_RF_PRE_FILTER_TEMP_SETTING | Temperature, Setting ¦ Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Temperature, Setting | Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Filter Replacement Fluid Temperature Setting | Temperature at which pre-filter replacement fluid is to be delivered to patient | Intradialytic | Parameter | Episodic | Numeric | XX.X | °C | Cel | MDC_DIM_DEGC | 1..1 | C7 |  | -- | -- | 0 to 99.9 | 41.0 | 3 |  | MMM | 258 |  | 27872 | 16936160
. . . MDC_HDIALY_RF_PRE_FILTER_TEMP | Temperature, Current ¦ Replacement Fluid, Pre-Filter ¦ Convective Pump ¦ Dialysis, Hemodialysis | Temperature, Current | Replacement Fluid, Pre-Filter | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Actual Pre-Filter Replacement Fluid Temperature | Temperature of pre-filter replacement fluid | Intradialytic | Status | Periodic | Numeric | XX.X | °C | Cel | MDC_DIM_DEGC | 1..1 | C7 |  | -- | -- | 0 to 99.9 | 39.9 | 3 |  | MMM | 2 |  | 27872 | 158944
. . . MDC_HDIALY_RF_CONV_CLEARANCE | Volume ¦ Replacement Fluid, Convective Urea Clearance ¦ Convective Pump ¦ Dialysis, Hemodialysis | Volume | Replacement Fluid, Convective Urea Clearance | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Convective Clearance - Urea | The product of the sieving coefficient and the total water flux. KX = QUF x SX where KX is Convective Clearance for solute "X", QUF is Ultrafiltration flow across membrane and SX is Membrane sieving coefficient for solute "X" | Intradialytic | Status | Periodic | Numeric | XXX.X | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 0..1 | O |  | -- | -- | 0 to 999.9 | 31.6 | 3 |  | MMM | 2 |  | 27880 | 158952
. . . MDC_EVT_HDIALY_RF_EXCESS_DELIV | LimitEvent ¦ Bolus Volume, ‹Excessive› ¦ Convective Pump ¦ Dialysis, Hemodialysis | LimitEvent | Bolus Volume, ‹Excessive› | Convective Pump | Dialysis, Hemodialysis | tech-high | HEMODIALYSIS | Replacement Fluid | Excessive Replacement Fluid Bolus Delivered | Notification that more replacement fluid has been delivered than intended | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1638 | 198246
. . . MDC_EVT_HDIALY_RF_INSUFF_DELIV | LimitEvent ¦ Bolus Volume, ‹Insufficient› ¦ Convective Pump ¦ Dialysis, Hemodialysis | LimitEvent | Bolus Volume, ‹Insufficient› | Convective Pump | Dialysis, Hemodialysis | tech-low | HEMODIALYSIS | Replacement Fluid | Insufficient Replacement Fluid Bolus Delivered | Notification that less replacement fluid has been delivered than intended | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1640 | 198248
. . . MDC_HDIALY_RF_PRE_POST_FLOW_RATIO_SETTING | Ratio ¦ Flow Rate (Pre and Post Replacement Fluid) ¦ Convective Pump ¦ Dialysis, Hemodialysis | Ratio | Flow Rate (Pre and Post Replacement Fluid) | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Pre-Post Replacement Fluid Flow Rate Ratio Setting | The ratio desired between the pre and post Replacement Fluid Flow Rate Settings | Intradialytic | Parameter | Episodic | Numeric | XX.X | N/A | -- | -- | 0..1 | O |  | -- | -- | 0 to 99.9 | 3.0 | 3 |  | MMM | 258 |  | 27884 | 16936172
. . . MDC_HDIALY_RF_BOLUS_RATE | Flow Rate ¦ Bolus ¦ Convective Pump ¦ Dialysis, Hemodialysis | Flow Rate | Bolus | Convective Pump | Dialysis, Hemodialysis | tech high low | HEMODIALYSIS | Replacement Fluid | Replacement Fluid Bolus Rate | User defined replacement fluid bolus rate. | Intradialytic | Parameter | Episodic | Numeric | XXXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 0..1 | C9 |  | -- | -- | 0 to 9999 | 100 | 3 |  | MMM | 2 |  | 27888 | 158960
. . . MDC_HDIALY_RF_BOLUS_VOL_SETTING | Volume, Setting ¦ Bolus ¦ Convective Pump ¦ Dialysis, Hemodialysis | Volume, Setting | Bolus | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Replacement Fluid Bolus Volume Setting | User defined amount of replacement fluid to be delivered in this single large dose | Intradialytic | Parameter | Episodic | Numeric | XXX | mL | mL | MDC_DIM_MILLI_L | 0..1 | C9 |  | -- | -- | 0 to 999 | 500 | 3 |  | MMM | 258 |  | 27892 | 16936180
. . . MDC_HDIALY_RF_BOLUS_VOL_DELIVERED | Volume, Delivered ¦ Bolus ¦ Convective Pump ¦ Dialysis, Hemodialysis | Volume, Delivered | Bolus | Convective Pump | Dialysis, Hemodialysis | tech high low | HEMODIALYSIS | Replacement Fluid | Replacement Fluid Bolus Volume Delivered | Amount of replacement fluid delivered in a single large dose. | Intradialytic | Parameter | Episodic | Numeric | XXX | mL | mL | MDC_DIM_MILLI_L | 0..1 | C9 |  | -- | -- | 0 to 999 | 489 | 3 |  | MMM | 2 |  | 27896 | 158968
. . . MDC_HDIALY_RF_SOURCE | Fluid Source (enumeration) ¦ Replacement Fluid ¦ Convective Pump ¦ Dialysis, Hemodialysis | Fluid Source (enumeration) | Replacement Fluid | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Replacement Fluid Source | The source of the replacement fluid. | Intradialytic | Status | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 0..1 | O | _TBL_09 – Fluid Source | BAG ONLINE | Bag – Replacement fluid Online – Replacement fluid prepared by dialysis machine | N/A | BAG | 3 |  | 1 | 2 |  | 27545 | 158617
. . . MDC_EVT_HDIALY_RF_BOLUS | StatusEvent ¦ Bolus, ‹Active› ¦ Convective Pump ¦ Dialysis, Hemodialysis | StatusEvent | Bolus, ‹Active› | Convective Pump | Dialysis, Hemodialysis |  | HEMODIALYSIS | Replacement Fluid | Replacement Fluid Bolus Event | True when an RF Bolus is occurring. | Intradialytic | Status | Episodic | Boolean | T/F | N/A | -- | -- | 0..1 | C9 |  | - | - | N/A | F | 3 | 0 | EVT | 3 |  | 1642 | 198250
. . MDC_DEV_HDIALY_SAFETY_SYSTEMS_CHAN | Subsystem ¦ ¦ Safety ¦ Dialysis, Hemodialysis | Subsystem |  | Safety | Dialysis, Hemodialysis |  |  |  | Hemodialysis safety system | Dialysis system safety events, not covered by other subsystems. | [8 empty] | 1..1 | M | [5 empty] | 3 |  | MVC (3) | 1 |  | 5427 | 70963
. . . MDC_EVT_HDIALY_SAFETY_ART_AIR_DETECT | ErrorEvent ¦ Arterial Blood, ‹Air In Line› ¦ Safety ¦ Dialysis, Hemodialysis | ErrorEvent | Arterial Blood, ‹Air In Line› | Safety | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Safety Systems | Arterial Air Detector | Notification that air is sensed in the arterial blood lines | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | M |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1644 | 198252
. . . MDC_EVT_HDIALY_SAFETY_DIALYSATE_AIR_DETECT | ErrorEvent ¦ Dialysate, ‹Air In Line› ¦ Safety ¦ Dialysis, Hemodialysis | ErrorEvent | Dialysate, ‹Air In Line› | Safety | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Safety Systems | Dialysate Air Detector | Notification that air is sensed in dialysate line | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | C6 |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1646 | 198254
. . . MDC_EVT_HDIALY_SAFETY_DIALYSATE_COMPOSITION | ErrorEvent ¦ Dialysate, Composition, ‹Incorrect› ¦ Safety ¦ Dialysis, Hemodialysis | ErrorEvent | Dialysate, Composition, ‹Incorrect› | Safety | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Safety Systems | Dialysis Composition | Notification that the dialyzing fluid composition is incorrect | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1648 | 198256
. . . MDC_EVT_HDIALY_SAFETY_SYSTEM_GENERAL | ErrorEvent ¦ Safety System, Non-Patient, ‹Error› ¦ Safety ¦ Dialysis, Hemodialysis | ErrorEvent | Safety System, Non-Patient, ‹Error› | Safety | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Safety Systems | General System | Notification of any dialysis machine malfunction that are not related to patient treatment | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | M |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1650 | 198258
. . . MDC_EVT_SELFTEST_FAILURE | ErrorEvent ¦ Selftest, ‹Error› ¦ Safety ¦ Dialysis, Hemodialysis | ErrorEvent | Selftest, ‹Error› | Safety | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Safety Systems | Self-Test | Notification that a self-test failed | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | M |  | -- | -- | N/A | F | 2 |  | EVT | 3 | 1608 | 1608 | 198216
. . . MDC_EVT_HDIALY_SAFETY_VEN_ACCESS | ErrorEvent ¦ Venous access, ‹Disconnection possible› ¦ Safety ¦ Dialysis, Hemodialysis | ErrorEvent | Venous access, ‹Disconnection possible› | Safety | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Safety Systems | Venous Access | Notification that the venous access may have become disconnected | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1652 | 198260
. . . MDC_EVT_HDIALY_SAFETY_VEN_AIR_DETECT | ErrorEvent ¦ Venous Blood, ‹Air In Line› ¦ Safety ¦ Dialysis, Hemodialysis | ErrorEvent | Venous Blood, ‹Air In Line› | Safety | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Safety Systems | Venous Air Detector | Notification that air is sensed in the venous blood lines | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | M |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1654 | 198262
. . . MDC_EVT_HDIALY_SAFETY_WETNESS_DETECT_ALERT | ErrorEvent ¦ Wetness detector, ‹Fluid detected› ¦ Safety ¦ Dialysis, Hemodialysis | ErrorEvent | Wetness detector, ‹Fluid detected› | Safety | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Safety Systems | Wetness Detector | Notification that the wetness detector has sensed fluid | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1656 | 198264
. . . MDC_EVT_HDIALY_SAFETY_WETNESS_DETECT_ERROR | ErrorEvent ¦ Wetness detector, ‹Error› ¦ Safety ¦ Dialysis, Hemodialysis | ErrorEvent | Wetness detector, ‹Error› | Safety | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Safety Systems | Wetness Detector Error | Notification of a detected error on the wetness detector device | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1658 | 198266
. . . MDC_ATTR_ALERT_ID_NUM | ErrorEvent (numeric) ¦ Safety Systen, ‹Notifications› ¦ Safety ¦ Dialysis, Hemodialysis | ErrorEvent (numeric) | Safety Systen, ‹Notifications› | Safety | Dialysis, Hemodialysis |  | HEMODIALYSIS | Safety Systems | Alert ID | The number associated with the current alert. If no alert is active then this object is not reported. | Intradialytic | Alert | Episodic | Numeric | XXXX | N/A | -- | -- | 0..1 | O |  | -- | -- | 0 to 9999 | 10 | 3 |  | 1 | 1 | 2953 | 2953 | 68489
. . . MDC_ATTR_ALERT_TEXT | ErrorEvent (string) ¦ Safety Systen, ‹Notifications› ¦ Safety ¦ Dialysis, Hemodialysis | ErrorEvent (string) | Safety Systen, ‹Notifications› | Safety | Dialysis, Hemodialysis |  | HEMODIALYSIS | Safety Systems | Alert Text | The text name of the current alert. If no alert is active then this object is not reported. IN most case this object will contain the text shown to the user. | Intradialytic | Alert | Episodic | String | N/A | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | Venous Air Detected | 1 |  | 1 | 1 | 3010 | 3010 | 68546
. . MDC_DEV_HDIALY_THERAPY_OUTCOMES_CHAN | Subsystem ¦ ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Subsystem |  | Therapy Outcomes | Dialysis, Hemodialysis |  |  |  | Hemodialysis therapy outcomes | Dialysis system patient treatment events, not covered by other subsystems. | [8 empty] | 1..1 | M | [5 empty] | 3 |  | MVC (3) | 1 |  | 5431 | 70967
. . . MDC_HDIALY_THERAPY_MASS_TRF_AREA_COEFF | Ratio ¦ Mass Transfer Coefficient x Surface Area of Dialyzer, (KoA) ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Ratio | Mass Transfer Coefficient x Surface Area of Dialyzer, (KoA) | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | KoA | Overall mass transfer coefficient multiplied by surface area of a dialyzer | Intradialytic | Status | Episodic | Numeric | XXXX | N/A | -- | -- | 0..1 | O |  | -- | -- | 0 to 9999 | 700 | 3 |  | MMM | 2 |  | 27900 | 158972
. . . MDC_HDIALY_THERAPY_ACCESS_FLOW | Flow Rate ¦ Patient Access ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Flow Rate | Patient Access | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Access Flow | Flow through a fistula or other conduit connecting an artery to a vein | Intradialytic | Status | Periodic | Numeric | ±XXXX | mL/min | mL/min | MDC_DIM_MILLI_L_PER_MIN | 0..1 | O |  | -- | -- | 0 to 9999 | 250 | 3 |  | MMM | 2 |  | 27904 | 158976
. . . MDC_HDIALY_THERAPY_RATIO_KT_OVER_V_GOAL | Ratio, Setting ¦ Urea Clearance ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Ratio, Setting | Urea Clearance | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Kt/V Goal | Target dialysis treatment adequacy measurement | Intradialytic | Parameter | Episodic | Numeric | X.XX | % | % | MDC_DIM_PERCENT | 0..1 | O |  | -- | -- | 0 to 9.99 | 1.21 | 3 |  | MMM | 2 |  | 27920 | 158992
. . . MDC_HDIALY_THERAPY_UREA_CLEARANCE_MEAN | Ratio, Mean ¦ Urea Clearance ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Ratio, Mean | Urea Clearance | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Mean Urea Clearance | Time weighted average of the individual effective conductivity clearance measurements | Intradialytic | Status | Episodic | Numeric | XXX | N/A | -- | -- | 0..1 | O |  | -- | -- | 0 to 999 | 70 | 3 |  | MMM (3) | 2 |  | 27947 | 159019
. . . MDC_HDIALY_THERAPY_UREA_CLEARANCE | Ratio ¦ Urea Clearance, Computed ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Ratio | Urea Clearance, Computed | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Urea Clearance | Calculated clearance based on the change in conductivity of the pre-dialyzer vs post dialyzer dialysate. | Intradialytic | Status | Episodic | Numeric | XXX | N/A | -- | -- | 0..1 | O |  | -- | -- | 0 to 999 | 196 | 3 |  | MMM | 2 |  | 27944 | 159016
. . . MDC_HDIALY_THERAPY_UREA_DISTRIBUTION_VOL_SETTING | Volume, Setting ¦ Urea ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Volume, Setting | Urea | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Urea Removed Setting | User setting for total urea to be removed during treatment. | All | Parameter | Episodic | Numeric | X.XX | L | L | MDC_DIM_L | 0..1 | O |  | -- | -- | 0 to 9999 | 100 | 8
. . . MDC_HDIALY_THERAPY_KT_DELIVERED | Volume, Removed ¦ Urea ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Volume, Removed | Urea | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Delivered Kt | Volume of fluid cleared of urea by dialysis treatment | Intradialytic | Status | Episodic | Numeric | X.XX | L | L | MDC_DIM_L | 0..1 | O |  | -- | -- | 0 to 9999.9 | 42.0 | 3 |  | MMM | 2 |  | 27912 | 158984
. . . MDC_HDIALY_THERAPY_RATIO_EKT_OVER_V_DELIVERED | Ratio ¦ Urea Clearance, Equilibrated, Delivered ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Ratio | Urea Clearance, Equilibrated, Delivered | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Delivered Equilibrated Kt/V | Fractional solute clearance for an intermittent treatment corrected to take account of solute redistribution following the treatment | Intradialytic | Status | Episodic | Numeric | X.XX | % | % | MDC_DIM_PERCENT | 0..1 | O |  | -- | -- | 0 to 9.99 | 1.1 | 3 |  | MMM | 2 |  | 27908 | 158980
. . . MDC_HDIALY_THERAPY_RATIO_SPKT_OVER_V_DELIVERED | Ratio ¦ Urea Clearance, Single Pool Model, Delivered ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Ratio | Urea Clearance, Single Pool Model, Delivered | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Delivered Single Pool Kt/V | Fractional solute clearance, calculated on the basis of solute concentration before and after an intermittent treatment, and on the basis of effluent dialysate volume and anthropometric estimate of volume of urea distribution for peritoneal dialysis | Intradialytic | Status | Episodic | Numeric | X.XX | % | % | MDC_DIM_PERCENT | 0..1 | O |  | -- | -- | 0 to 9.99 | 1.1 | 3 |  | MMM | 2 |  | 27916 | 158988
. . . MDC_HDIALY_THERAPY_RATIO_SPKT_OVER_V_PROJECTED | Ratio ¦ Urea Clearance, Single Pool Model, Projected ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Ratio | Urea Clearance, Single Pool Model, Projected | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Projected Single Pool Kt/V | Fractional solute clearance, calculated on the basis of solute concentration before and after an intermittent treatment, and on the basis of effluent dialysate volume and anthropometric estimate of volume of urea distribution for peritoneal dialysis, projected on the basis of data available before the treatment occurs or before the treatment is completed | Intradialytic | Status | Episodic | Numeric | X.XX | % | % | MDC_DIM_PERCENT | 0..1 | O |  | -- | -- | 0 to 9.99 | 1.1 | 3 |  | MMM | 2 |  | 27940 | 159012
. . . MDC_EVT_HDIALY_THERAPY_PAT_TX_GENERAL | StatusEvent ¦ Treatment, Non-specific alarm condition, ‹Error› ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | StatusEvent | Treatment, Non-specific alarm condition, ‹Error› | Therapy Outcomes | Dialysis, Hemodialysis | phys | HEMODIALYSIS | Therapy Outcomes | General Patient Treatment | Notification that the machine has detected an alarm condition not covered by other alarms that is related to patient treatment. | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1660 | 198268
. . . MDC_HDIALY_THERAPY_PCT_RECIRC | Ratio ¦ Blood Flow, in Access Line (percentage dialyzed blood returned immediately back to the dialysis machine instead of entering back into the patient's vasculature) ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Ratio | Blood Flow, in Access Line (percentage dialyzed blood returned immediately back to the dialysis machine instead of entering back into the patient's vasculature) | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Percent Recirculation | Percentage of the blood flow in the access line that is coming directly from the venous line without going through the patient’s body. | Intradialytic | Status | Periodic | Numeric | ±XXX.X | % | % | MDC_DIM_PERCENT | 0..1 | O |  | -- | -- | 0 to 999.9 | 20.0 | 3 |  | MMM | 2 |  | 27932 | 159004
. . . MDC_HDIALY_THERAPY_PLASMA_NA_CONC | Concentration ¦ Sodium (NA), Patient Plasma ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Concentration | Sodium (NA), Patient Plasma | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Plasma Sodium Concentration | Sodium ion concentration measured in patient plasma. | Intradialytic | Status | Episodic | Numeric | XXX | mmol/L | mmol/L | MDC_DIM_MILLI_MOLE_PER_L | 0..1 | O |  | -- | -- | 0 to 999 | 140 | 3 |  | MMM | 2 |  | 27936 | 159008
. . . MDC_HDIALY_THERAPY_COMPLETE_METHOD | Method (enumeration) ¦ Treatment Completion ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Method (enumeration) | Treatment Completion | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Treatment Completion Method | Parameter used to determine when treatment has been complete | Intradialytic | Status | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 1..1 | M | _TBL_12 – Completion Method | CT AT UF KTV USER | Clock Time Actual Treatment Time UF Removed Kt/V The treatment runs until stopped by the user. | N/A | UF | 3 |  | 1 | 2 |  | 27546 | 158618
. . . MDC_EVT_HDIALY_THERAPY_TX_END_TIME | StatusEvent ¦ Treament Time, ‹Achieved› ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | StatusEvent | Treament Time, ‹Achieved› | Therapy Outcomes | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Therapy Outcomes | Treatment Time Ended | Notification that target time of treatment has been met | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1662 | 198270
. . . MDC_HDIALY_THERAPY_BODY_START_WT | Weight ¦ Body, Start of Treatment ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Weight | Body, Start of Treatment | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Patient Start Weight | Weight of patient at beginning of treatment | All | Parameter | Episodic | Numeric | XXX.X | Kg | kg | MDC_DIM_KILO_G | 0..1 | O |  | -- | -- | 0 to 999.9 | 75.9 | 3 |  | MMM | 2 |  | 27928 | 159000
. . . MDC_HDIALY_THERAPY_BODY_END_WT | Weight ¦ Body, End of Treatment ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Weight | Body, End of Treatment | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Patient End Weight | Weight of patient at end of treatment | All | Parameter | Episodic | Numeric | XXX.X | Kg | kg | MDC_DIM_KILO_G | 0..1 | O |  | -- | -- | 0 to 999.9 | 75.9 | 3 |  | MMM | 2 |  | 27948 | 159020
. . . MDC_HDIALY_THERAPY_BODY_END_WT_SETTING | Weight, Setting ¦ Body, End of Treatment ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Weight, Setting | Body, End of Treatment | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Patient End Weight Target | Desired weight of patient at end of treatment | All | Parameter | Episodic | Numeric | XXX.X | Kg | kg | MDC_DIM_KILO_G | 0..1 | O |  | -- | -- | 0 to 999.9 | 75.0 | 8
. . . MDC_HDIALY_THERAPY_END_TIME_SETTING | Time, Setting ¦ Therapy End ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Time, Setting | Therapy End | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Treatment End Time | The clock time that the therapy ends when completion method is Clock Time, | All | Parameter | Episodic | Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | -- | 0..1 | O |  | -- | -- |  |  | 8
. . . MDC_HDIALY_THERAPY_TX_TIME_SETTING | Duration, Setting ¦ Therapy End ¦ Therapy Outcomes ¦ Dialysis, Hemodialysis | Duration, Setting | Therapy End | Therapy Outcomes | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Treatment Time | The duration of the therapy when completion method is Tx Time. | All | Parameter | Episodic | Numeric | XXXX | minutes | min | MDC_DIM_MIN | 0..1 | O |  | -- | -- | 0 to 9999 | 300 | 8
. . MDC_DEV_HDIALY_UF_CHAN | Subsystem ¦ ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | Subsystem |  | Ultrafiltration | Dialysis, Hemodialysis |  |  |  | Hemodialysis ultrafiltration | The observations, settings and events related to the removal of fluid from the patient (aka. ultrafiltration). | [8 empty] | 1..1 | M | [5 empty] | 3 |  | MVC (3) | 1 |  | 5435 | 70971
. . . MDC_HDIALY_UF_MODE | Pump Mode (enumeration) ¦ Ultrafiltration ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | Pump Mode (enumeration) | Ultrafiltration | Ultrafiltration | Dialysis, Hemodialysis |  | HEMODIALYSIS | Ultrafiltration System | Ultrafiltration Mode | Process used to remove fluid volume from the patient | Intradialytic | Parameter | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 1..1 | M | _TBL_13 – Ultrafiltration Mode | NONE CONST-WT PRO-WT CONST-WOT PRO-WOT | None UF removed at a constant rate with a target volume UF removed according to a programmed profile with a Target Volume UF removal at a constant rate without a target Volume UF removed according to programmed profile without a target volume. | N/A | CUF | 3 |  | 1 | 2 |  | 27547 | 158619
. . . .Profile Parameters | [9 empty] | Contains the ultrafiltration profile parameters. | [8 empty] | 0..1 | C24
. . . MDC_HDIALY_UF_RATE_SETTING | Flow Rate, Setting ¦ Ultrafiltration ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | Flow Rate, Setting | Ultrafiltration | Ultrafiltration | Dialysis, Hemodialysis |  | HEMODIALYSIS | Ultrafiltration System | Ultrafiltration Rate | Programed rate at which fluid is removed from the patient. | Intradialytic | Status | Periodic | Numeric | XXXX | mL/h | mL/h | MDC_DIM_MILLI_L_PER_HR | 1..1 | M |  |  |  | 0 to 9999 | 100 | 3 |  | MMM | 258 |  | 27964 | 16936252
. . . MDC_HDIALY_UF_RATE | Flow Rate, Current ¦ Ultrafiltration ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | Flow Rate, Current | Ultrafiltration | Ultrafiltration | Dialysis, Hemodialysis |  | HEMODIALYSIS | Ultrafiltration System | Ultrafiltration Rate | Current rate at which fluid is removed from the patient | Intradialytic | Status | Periodic | Numeric | XXXX | mL/h | mL/h | MDC_DIM_MILLI_L_PER_HR | 1..1 | M |  | -- | -- | 0 to 9999 | 100 | 3 |  | MMM | 2 |  | 27964 | 159036
. . . MDC_EVT_HDIALY_UF_LO | LimitEvent ¦ Volume Removed, ‹Insufficient› ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | LimitEvent | Volume Removed, ‹Insufficient› | Ultrafiltration | Dialysis, Hemodialysis | tech-low | HEMODIALYSIS | Ultrafiltration System | Insufficient Ultrafiltration | Notification that the ultrafiltration system is not removing as much fluid as intended | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1664 | 198272
. . . MDC_EVT_HDIALY_UF_NEG | LimitEvent ¦ Effective Rate, ‹Negative› ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | LimitEvent | Effective Rate, ‹Negative› | Ultrafiltration | Dialysis, Hemodialysis | tech-low | HEMODIALYSIS | Ultrafiltration System | Negative Ultrafiltration | Notification that the ultrafiltration system has resulted in a negative ultrafiltration rate | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1666 | 198274
. . . MDC_HDIALY_UF_TIME_TO_TARGET | Duration ¦ Time Until Target ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | Duration | Time Until Target | Ultrafiltration | Dialysis, Hemodialysis |  | HEMODIALYSIS | Ultrafiltration System | Remaining Ultrafiltration Time | Time remaining until target volume to remove will have been achieved. | Intradialytic | Status | Periodic | Numeric | XXXX | minutes | min | MDC_DIM_MIN | 0..1 | O |  | -- | -- | 0 to 9999 | 55 | 3 |  | MMM | 2 |  | 27952 | 159024
. . . MDC_HDIALY_UF_TARGET_VOL_TO_REMOVE | Volume, Setting ¦ Ultrafiltration ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | Volume, Setting | Ultrafiltration | Ultrafiltration | Dialysis, Hemodialysis |  | HEMODIALYSIS | Ultrafiltration System | Target Volume To Remove | Planned fluid removal | Intradialytic | Parameter | Episodic | Numeric | XXXX | mL | mL | MDC_DIM_MILLI_L | 1..1 | C11 |  | -- | -- | 0 to 9999 | 2000 | 3 |  | MMM | 2 |  | 27956 | 159028
. . . MDC_HDIALY_UF_ACTUAL_REMOVED_VOL | Volume, Removed ¦ Ultrafiltration ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | Volume, Removed | Ultrafiltration | Ultrafiltration | Dialysis, Hemodialysis |  | HEMODIALYSIS | Ultrafiltration System | Total Fluid Volume Removed | Measured fluid removed from patient | Intradialytic | Status | Periodic | Numeric | XXXX | mL | mL | MDC_DIM_MILLI_L | 1..1 | M |  | -- | -- | 0 to 9999 | 555 | 3 |  | MMM | 2 |  | 27960 | 159032
. . . MDC_EVT_HDIALY_UF_RATE_RANGE | LimitEvent ¦ Effective Rate, ‹OutOfRange› ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | LimitEvent | Effective Rate, ‹OutOfRange› | Ultrafiltration | Dialysis, Hemodialysis | tech-range | HEMODIALYSIS | Ultrafiltration System | Ultrafiltration Rate Out of Range | Notification that the ultrafiltration rate has deviated from the intended ultrafiltration rate more than is safe. Specified another way, MDCX_HDIALY_UF_RATE is not between MDCX_HDIALY_UF_RATE_SETTING_MIN and MDCX_HDIALY_UF_RATE_SETTING_MAX | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | M |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1668 | 198276
. . . MDC_EVT_HDIALY_UF_GOAL_MET | StatusEvent ¦ Target Volume, ‹Achieved› ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | StatusEvent | Target Volume, ‹Achieved› | Ultrafiltration | Dialysis, Hemodialysis | tech | HEMODIALYSIS | Ultrafiltration System | Ultrafiltration Goal Met | Notification that target ultrafiltration volume has been met | Intradialytic | Alert | Episodic | Boolean | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | 3 | 0 | EVT | 3 |  | 1670 | 198278
. . . MDC_HDIALY_UF_RATE_LIMIT_HIGH_SETTING | Flow Rate, Setting ¦ Ultrafiltration, Maximum ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | Flow Rate, Setting | Ultrafiltration, Maximum | Ultrafiltration | Dialysis, Hemodialysis |  | HEMODIALYSIS | Ultrafiltration System | Ultrafiltration Flow Rate Maximum Limit | Maximum ultrafiltration rate allowed without an alarm generation | Intradialytic | Parameter | Episodic | Numeric | XXXX | mL/h | mL/h | MDC_DIM_MILLI_L_PER_HR | 0..1 | O |  | -- | -- | 0 to 9999 | 100 | 3 |  | 1 | 258 |  | 27969 | 16936257
. . . MDC_HDIALY_UF_RATE_LIMIT_LOW_SETTING | Flow Rate, Setting ¦ Ultrafiltration, Minimum ¦ Ultrafiltration ¦ Dialysis, Hemodialysis | Flow Rate, Setting | Ultrafiltration, Minimum | Ultrafiltration | Dialysis, Hemodialysis |  | HEMODIALYSIS | Ultrafiltration System | Ultrafiltration Flow Rate Minimum Limit | Minimum ultrafiltration rate allowed without an alarm generation | Intradialytic | Parameter | Episodic | Numeric | XXXX | mL/h | mL/h | MDC_DIM_MILLI_L_PER_HR | 0..1 | O |  | -- | -- | 0 to 9999 | 100 | 3 |  | 1 | 258 |  | 27971 | 16936259
. MDC_DEV_PRESS_BLD_NONINV_VMD | Device ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Device | Pressure, noninvasive | Blood | CVS |  |  |  | Non-invasive blood pressure | Instrument for the non-invasive measurement of blood pressure. | [8 empty] | 0..1 | O | [5 empty] | P |  | MVC (2) | 1 | 5150 | 5150 | 70686
. . MDC_DEV_PRESS_BLD_NONINV_CHAN | Subsystem ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Subsystem | Pressure, noninvasive | Blood | CVS | [13 empty] | 0..1 | O | [5 empty] | P |  | MVC (3) | 1 | 5151 | 5151 | 70687
. . . MDC_ATTR_TIME_PD_MSMT | Duration ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Duration | Pressure, noninvasive | Blood | CVS |  | HEMODIALYSIS | Blood Pressure Monitor | Blood Pressure Measurement Interval Setting | Time allowed to elapse between automatic blood pressure measurements. A value of zero indicates that automatic measurements are not being taken. | All | Parameter | Episodic | Numeric | XXX | minutes | min | MDC_DIM_MIN | 0..1 | O |  | -- | -- | 0 to 999 | 30 | P |  | 1 | 1 | 2443 | 2443 | 67979
. . . MDC_ATTR_PT_BODY_POSN | Body Position (enumeration) ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Body Position (enumeration) | Pressure, noninvasive | Blood | CVS |  | HEMODIALYSIS | Blood Pressure Monitor | Patient body position | Patient body position at the time of the {blood pressure} measurement. | All | Status | Episodic | Enumeration | Alphanumeric | N/A | -- | -- | 0..1 | O | _TBL_04 – Patient Position | SITTING STANDING SUPINE | Sitting Standing Supine | N/A | SUPINE | 3 |  | 1 | 1 | 2599 | 2599 | 68135
. . . MDC_PRESS_BLD_NONINV_DIA | Pressure ¦ Noninvasive, Diastolic ¦ Blood ¦ CVS | Pressure | Noninvasive, Diastolic | Blood | CVS | phys high low thr | HEMODIALYSIS | Blood Pressure Monitor | Diastolic Pressure | Minimum arterial pressure during relaxation and dilatation of the ventricles of the heart when the ventricles fill with blood | All | Status | Episodic | Numeric | XXX | mmHg | mm[Hg] | MDC_DIM_MMHG | 0..1 | O |  | -- | -- | 0 to 999 | 80 | P |  | SDM (2) | 2 | 18950 | 18950 | 150022
. . . MDC_PULS_RATE_NON_INV | Rate ¦ Noninvasive, Pulse ¦ Blood ¦ CVS | Rate | Noninvasive, Pulse | Blood | CVS | phys high low thr | HEMODIALYSIS | Blood Pressure Monitor | Heart Rate | Number of heart beats per minute | All | Status | Episodic | Numeric | XXX | BPM | {beats}/min | MDC_DIM_BEAT_PER_MIN | 0..1 | O |  | -- | -- | 0 to 999 | 70 | P |  | RCE (2),, | 2 | 18474 | 18474 | 149546
. . . MDC_PRESS_BLD_NONINV_MEAN | Pressure ¦ Noninvasive, Mean ¦ Blood ¦ CVS | Pressure | Noninvasive, Mean | Blood | CVS |  | HEMODIALYSIS | Blood Pressure Monitor | Mean Arterial Pressure | Average pressure in a patient's arteries during one cardiac cycle | All | Status | Episodic | Numeric | XXX | mmHg | mm[Hg] | MDC_DIM_MMHG | 0..1 | O |  | -- | -- | 0 to 999 | 100 | P |  | SDM (3) | 2 | 18951 | 18951 | 150023
. . . MDC_PRESS_BLD_NONINV_SYS | Pressure ¦ Noninvasive, Systolic ¦ Blood ¦ CVS | Pressure | Noninvasive, Systolic | Blood | CVS | phys high low thr | HEMODIALYSIS | Blood Pressure Monitor | Systolic Pressure | Maximum arterial pressure during contraction of the left ventricle of the heart | All | Status | Episodic | Numeric | XXX | mmHg | mm[Hg] | MDC_DIM_MMHG | 0..1 | O |  | -- | -- | 0 to 999 | 120 | P |  | SDM (1) | 2 | 18949 | 18949 | 150021
. MDC_DEV_ANALY_SAT_O2_VMD | Device ¦ Concentration [Sat] ¦ Blood ¦ Blood Chemistry | Device | Concentration [Sat] | Blood | Blood Chemistry |  |  |  | SpO2 monitor | Instrument that derives the % of arterial O2 saturation and pulse rate. | [8 empty] | 0..1 | O | [5 empty] | P |  | MVC (2) | 1 | 4106 | 4106 | 69642
. . MDC_DEV_ANALY_SAT_O2_CHAN | Subsystem ¦ Concentration [Sat] ¦ Blood ¦ Blood Chemistry | Subsystem | Concentration [Sat] | Blood | Blood Chemistry | [13 empty] | 0..1 | O | [5 empty] | P |  | MVC (3) | 1 | 4107 | 4107 | 69643
. . . MDC_PULS_OXIM_SAT_O2 | Concentration ¦ Saturation, Oxygen ¦ Blood ¦ Blood Chemistry | Concentration | Saturation, Oxygen | Blood | Blood Chemistry | phys high low thr | HEMODIALYSIS | Pulse Oximeter Monitor | SpO2 Oxygen Saturation | Hemoglobin oxygen saturation in patient blood measured by pulse oximetry | Intradialytic | Status | Periodic | Numeric | XXX | % | % | MDC_DIM_PERCENT | 0..1 | O |  | -- | -- | 0 to 999 | 98 | P |  | MMM | 2 | 19384 | 19384 | 150456
. . . MDC_PULS_OXIM_PULS_RATE | Rate ¦ Oximetry ¦ Pulse ¦ Blood, CVS | Rate | Oximetry | Pulse | Blood, CVS | phys high low thr | HEMODIALYSIS | Pulse Oximeter Monitor | SpO2 Pulse Rate | Number of heart beats per minute measured by pulse oximetry | Intradialytic | Status | Periodic | Numeric | XXX | BPM | {beats}/min | MDC_DIM_BEAT_PER_MIN | 0..1 | O |  | -- | -- | 0 to 999 | 67 | P |  | RCE (2) | 2 | 18458 | 18458 | 149530
. . . MDC_EVT_ERR | ErrorEvent ¦ Event, ‹Error› ¦ Processing ¦ Device | ErrorEvent | Event, ‹Error› | Processing | Device | tech | HEMODIALYSIS | Safety Systems | Pulse Oximeter Error | Notification of a detected error on the pulse oximeter device | Intradialytic | Alert | Episodic | Bool | T / F | N/A | -- | -- | 0..1 | O |  | -- | -- | N/A | F | P |  | EVT | 3 | 30 | 30 | 196638
. MDC_DEV_BLOOD_CHEM_VMD | Device ¦ Multiparameter, Multimethod ¦ Blood Chemistry ¦ Dialysis, Hemodialysis | Device | Multiparameter, Multimethod | Blood Chemistry | Dialysis, Hemodialysis |  |  |  | Blood Chemistry Monitor | Blood chemistry monitor that may include multiple analytes and multiple aquisition, analysis and data capture methods (e.g. optical, ultrasound) and may include manually entered measurements. | [8 empty] | 0..1 | O | [5 empty] | 3 |  | MVC (2) | 1 |  | 5438 | 70974
. . MDC_DEV_BLOOD_CHEM_CHAN | Subsystem ¦ Multiparameter, Multimethod ¦ Blood Chemistry ¦ Dialysis, Hemodialysis | Subsystem | Multiparameter, Multimethod | Blood Chemistry | Dialysis, Hemodialysis | [13 empty] | 0..1 | O | [5 empty] | 3 |  | MVC (3) | 1 |  | 5439 | 70975
. . . MDC_HDIALY_PLASMA_VOL_MARKER | Indication (enumeration) ¦ Symptom or Intervention associated with the relative blood volume percent reading ¦ Blood Chemistry ¦ Dialysis, Hemodialysis | Indication (enumeration) | Symptom or Intervention associated with the relative blood volume percent reading | Blood Chemistry | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Plasma Volume Marker | An indicator applied to the relative blood volume percent reading to designate an event | Intradialytic | Status | Periodic | Enumeration | Alphanumeric | N/A | -- | -- | 0..1 | O | _TBL_10 – Plasma Volume Marker | NONE SYMP INTER | None – No selection made at time of reading/reporting Symptomatic – Noting a symptom for continue monitoring. Intervention – Action was taken or a machine setting was changed. | N/A | NONE | 3 |  | 1 | 2 |  | 27548 | 158620
. . . MDC_HDIALY_PLASMA_VOL_PROFILE | Profile (enumeration) ¦ Rate of change for relative blood volume ¦ Blood Chemistry ¦ Dialysis, Hemodialysis | Profile (enumeration) | Rate of change for relative blood volume | Blood Chemistry | Dialysis, Hemodialysis |  | HEMODIALYSIS | Therapy Outcomes | Plasma Volume Profile | The rate of change for relative blood volume. | Intradialytic | Status | Periodic | Enumeration | Alphanumeric | N/A | -- | -- | 0..1 | O | _TBL_11 – Plasma Volume Profile | FLAT GRAD STEEP | Change less than or equal to 3% per hour Change greater than 3% and less than or equal to 6.5% per hour Change greater than 6.5% per hour. | N/A | FLAT | 3 |  | 1 | 2 |  | 27549 | 158621
. . . MDC_CONC_HCT_GEN | Concentration ¦ Total, HCT ¦ General ¦ Fluid Chemistry | Concentration | Total, HCT | General | Fluid Chemistry | phys low high thr | HEMODIALYSIS | Therapy Outcomes | Hematocrit | 100*(sample red blood cell volume/blood sample volume) | Intradialytic | Status | Periodic | Numeric | XX.X | % | %{vol} | MDC_DIM_VOL_PERCENT | 0..1 | O |  | -- | -- | 0 to 99.9 | 45 | P |  | MMM | 2 | 29060 | 29060 | 160132
. . . MDC_CONC_HB_GEN | Concentration ¦ Total, HgB ¦ General ¦ Fluid Chemistry | Concentration | Total, HgB | General | Fluid Chemistry | phys low high thr | HEMODIALYSIS | Therapy Outcomes | Hemoglobin | Concentration of hemoglobin in a sample expressed as mass/volume | Intradialytic | Status | Periodic | Numeric | ±XXX.X | g/dL | g/dL | MDC_DIM_G_PER_DL | 0..1 | O |  | -- | -- | 0 to 999.9 | 13.6 | P |  | MMM | 2 | 29048 | 29048 | 160120
. . . MDC_HDIALY_REL_BLOOD_VOL | Ratio ¦ Blood Volume Change since the start of treatment ¦ Blood Chemistry ¦ Dialysis, Hemodialysis | Ratio | Blood Volume Change since the start of treatment | Blood Chemistry | Dialysis, Hemodialysis | phys low thr | HEMODIALYSIS | Therapy Outcomes | Relative Blood Volume | The percent change in blood volume since the start of the treatment. The BV change value is calculated from the following equation: ΔBV% = [(H0/H1) - 1] x 100 Where H0 = initial Hct or Hb H1 = current Hct or Hb | Intradialytic | Status | Periodic | Numeric | ±XXX.X | % | % | MDC_DIM_PERCENT | 0..1 | O |  | -- | -- | ±XXX | -15 | 3 |  | MMM | 2 |  | 27972 | 159044
. . . MDC_SAT_O2 | Concentration ¦ Saturation, Oxygen ¦ Blood ¦ Blood Chemistry | Concentration | Saturation, Oxygen | Blood | Blood Chemistry | phys low high thr | HEMODIALYSIS | Therapy Outcomes | SpO2 Oxygen Saturation | Hemoglobin oxygen saturation in patient blood | Intradialytic | Status | Periodic | Numeric | XXX | % | % | MDC_DIM_PERCENT | 0..1 | O |  | -- | -- | 0 to 999 | 98 | P |  | MMM | 2 | 19244 | 19244 | 150316

## DeviceAlarms
source | event | AlertType | Dialysis Common Name | Definition | Mandatory (M), Optional (O), Conditional (C, mandatory condition)
MDCX_DEV_HDIALY_MACHINE_MDS
MDCX_DEV_HDIALY_VMD
MDCX_DEV_HDIALY_MACH_CONFIG_CHAN
MDCX_DEV_HDIALY_ANTICOAG_PUMP_CHAN
MDCX_HDIALY_ANTICOAG_PUMP_CHAN | MDCX_EVT_HDIALY_ANTICOAG_STOP | tech | Anticoagulant Pump Stop | Notification that the anticoagulant pump has stopped | O
MDCX_HDIALY_ANTICOAG_PUMP_CHAN | MDCX_EVT_HDIALY_ANTICOAG_SYRINGE_EMPTY | tech | Syringe Empty | Notification that the syringe plunger has reached the end of its functional range | O
MDCX_HDIALY_ANTICOAG_PUMP_CHAN | MDCX_EVT_HDIALY_ANTICOAG_SYRINGE_SIZE | tech | Wrong Syringe Size | Notification that the size of the syringe is not the size defined for anticoagulant administration | O
MDCX_DEV_HDIALY_BLOOD_PUMP_CHAN
MDCX_HDIALY_BLD_PRESS_ART | MDC_EVT_HI | tech high thr | Arterial Pressure High | Notification that the arterial pressure has exceeded maximum setting | M
MDCX_HDIALY_BLD_PRESS_ART | MDC_EVT_LO | tech low thr | Arterial Pressure Low | Notification that the arterial pressure has fallen below the minimum setting | M
MDCX_HDIALY_BLOOD_PUMP_CHAN | MDCX_EVT_HDIALY_BLD_PUMP_STOP | tech | Blood Pump Stop | Notification that the blood pump has stopped | M
MDCX_HDIALY_BLD_PUMP_PRESS_ART_POST_PUMP | MDC_EVT_HI | tech high thr | Post Pump Arterial Pressure High | Notification that the post-pump arterial pressure is above limit. | O
MDCX_HDIALY_BLD_PUMP_PRESS_ART_POST_PUMP | MDC_EVT_LO | tech low thr | Post Pump Arterial Pressure Low | Notification that the post-pump arterial pressure is below limit. | O
MDCX_HDIALY_BLD_PUMP_SINGLE_NEEDLE_PRESS | MDC_EVT_HI | tech high thr | Needle Pressure High | Notification that the needle pressure is above limit. | O
MDCX_HDIALY_BLD_PUMP_SINGLE_NEEDLE_PRESS | MDC_EVT_LO | tech low thr | Needle Pressure Low | Notification that the needle pressure is below limit. | O
MDCX_HDIALY_BLD_PUMP_PRESS_VEN | MDC_EVT_HI | tech high thr | Venous Pressure High | Notification that the venous pressure has exceeded maximum setting | M
MDCX_HDIALY_BLD_PUMP_PRESS_VEN | MDC_EVT_LO | tech low thr | Venous Pressure Low | Notification that the venous pressure has fallen below the minimum setting | M
MDCX_DEV_HDIALY_FLUID_CHAN
MDCX_HDIALY_DIALYSATE_COND | MDC_EVT_HI | tech high | Conductivity High | Notification that the conductivity of the dialysate has exceeded the maximum setting | O
MDCX_HDIALY_DIALYSATE_COND | MDC_EVT_LO | tech low | Conductivity Low | Notification that the conductivity of the dialysate has fallen below the minimum setting | O
MDCX_HDIALY_DIALYSATE_FLOW_RATE | MDC_EVT_LO | tech low | Dialysate Flow Low | Notification that the dialysate flow is below the dialysate flow rate minimum setting | O
MDCX_HDIALY_DIALYSATE_TEMP | MDC_EVT_LO | tech low | Dialysate Temperature Low | Notification that the dialysate temperate has fallen below the minimum temperate setting | C (if mode is HD or HDF)
MDCX_HDIALY_DIALYSATE_TEMP | MDC_EVT_HI | tech high | Dialysate Temperature High | Notification that the dialysate temperate exceeds the maximum temperate setting | C (if mode is HD or HDF)
MDCX_HDIALY_FLUID_CHAN | MDCX_EVT_HDIALY_BLOOD_LEAK | tech | Blood Leak | Notification that blood has been detected in the dialysate fluid | M
MDCX_DEV_HDIALY_FILTER_CHAN
MDCX_HDIALY_FILTER_TRANSMEMBRANE_PRESS | MDC_EVT_LO | tech low | Transmembrane Pressure Low | Notification that transmembrane pressure has fallen below the minimum setting | M
MDCX_HDIALY_FILTER_TRANSMEMBRANE_PRESS | MDC_EVT_HI | tech high | Transmembrane Pressure High | Notification that transmembrane pressure has exceeded the maximum setting | M
MDCX_DEV_HDIALY_CONVECTIVE_CHAN
MDCX_HDIALY_CONVECTIVE_CHAN | MDCX_EVT_HDIALY_RF_EXCESS_DELIV | tech-high | Excessive Replacement Fluid Bolus Delivered | Notification that more replacement fluid has been delivered than intended | O
MDCX_HDIALY_CONVECTIVE_CHAN | MDCX_EVT_HDIALY_REPLACE_FLUID_INSUFF_DELIV | tech-low | Insufficient Replacement Fluid Bolus Delivered | Notification that less replacement fluid has been delivered than intended | O
MDCX_HDIALY_RF_BOLUS_RATE | MDC_EVT_HI | tech high | Replacement Fluid Bolus Rate High | Notification that the replacement fluid bolus is delivering more fluid than intended | O
MDCX_HDIALY_RF_BOLUS_RATE | MDC_EVT_LO | tech low | Replacement Fluid Bolus Rate Low | Notification that the replacement fluid bolus is delivering less than the intended. | O
MDCX_HDIALY_RF_BOLUS_VOL_DELIVERED | MDC_EVT_HI | tech high | Replacement Fluid Bolus Volume High | Notification that the replacement fluid bolus delivered more fluid than intended | O
MDCX_HDIALY_RF_BOLUS_VOL_DELIVERED | MDC_EVT_LO | tech low | Replacement Fluid Bolus Volume Low | Notification that the replacement fluid bolus delivered less fluid than intended. | O
MDCX_DEV_HDIALY_SAFETY_SYSTEMS_CHAN
MDCX_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_ART_AIR_DETECT | tech | Arterial Air Detector | Notification that air is sensed in the arterial blood lines | M
MDCX_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_DIALYSATE_AIR_DETECT | tech | Dialysate Air Detector | Notification that air is sensed in dialysate line | C (if mode is HD or HDF)
MDCX_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_DIALYSATE_COMPOSITION | tech | Dialysis Composition | Notification that the dialyzing fluid composition is incorrect | O
MDCX_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_SYSTEM_GENERAL | tech | General System | Notification of any dialysis machine malfunction that are not related to patient treatment | M
MDCX_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_SELF_TEST | tech | Self-Test | Notification that a self-test failed | M
MDCX_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_VEN_ACCESS | tech | Venous Access | Notification that the venous access may have become disconnected | O
MDCX_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_VEN_AIR_DETECT | tech | Venous Air Detector | Notification that air is sensed in the venous blood lines | M
MDCX_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_WETNESS_DETECT_ALERT | tech | Wetness Detector | Notification that the wetness detector has sensed fluid | O
MDCX_HDIALY_SAFETY_SYSTEMS_CHAN | MDCX_EVT_HDIALY_SAFETY_WETNESS_DETECT_ERROR | tech | Wetness Detector Error | Notification of a detected error on the wetness detector device | O
MDCX_DEV_HDIALY_THERAPY_OUTCOMES_CHAN
MDCX_HDIALY_THERAPY_OUTCOMES_CHAN | MDCX_EVT_HDIALY_THERAPY_PAT_TX_GENERAL | phys | General Patient Treatment | Notification that the machine has detected an alarm condition not covered by other alarms that is related to patient treatment. | O
MDCX_HDIALY_THERAPY_OUTCOMES_CHAN | MDCX_EVT_HDIALY_THERAPY_TX_END_TIME | tech | Treatment Time Ended | Notification that target time of treatment has been met | O
MDCX_DEV_HDIALY_UF_CHAN
MDCX_HDIALY_UF_CHAN | MDCX_EVT_HDIALY_UF_LO | tech low | Insufficient Ultrafiltration | Notification that the ultrafiltration system is not removing as much fluid as intended | O
MDCX_HDIALY_UF_CHAN | MDCX_EVT_HDIALY_UF_NEG | tech low | Negative Ultrafiltration | Notification that the ultrafiltration system has resulted in a negative ultrafiltration rate | O
MDCX_HDIALY_UF_CHAN | MDCX_EVT_HDIALY_UF_GOAL_MET | tech | Ultrafiltration Goal Met | Notification that target ultrafiltration volume has been met | O
MDCX_HDIALY_UF_CHAN | MDCX_EVT_HDIALY_UF_RATE_RANGE | tech range | Ultrafiltration Rate Out of Range | Notification that the ultrafiltration rate has deviated from the intended ultrafiltration rate more than is safe | M
MDC_DEV_PRESS_BLD_NONINV_VMD
MDC_DEV_PRESS_BLD_NONINV_CHAN
MDC_PRESS_BLD_NONINV_DIA | MDC_EVT_HI | phys high thr | Diastolic High | Notification that diastolic pressure exceeded the diastolic maximum setting | O
MDC_PRESS_BLD_NONINV_DIA | MDC_EVT_LOW | phys low thr | Diastolic Low | Notification that diastolic pressure fell below the diastolic minimum setting | O
MDC_PULS_RATE_NON_INV | MDC_EVT_HI | phys high thr | Heart Rate High | Heart rate exceeded heart rate maximum setting | O
MDC_PULS_RATE_NON_INV | MDC_EVT_LOW | phys low thr | Heart Rate Low | Heart rate fell below heart rate minimum setting | O
MDC_PRESS_BLD_NONINV_SYS | MDC_EVT_HI | phys high thr | Systolic High | Notification that systolic pressure exceeded the systolic maximum setting | O
MDC_PRESS_BLD_NONINV_SYS | MDC_EVT_LOW | phys low thr | Systolic Low | Notification that systolic pressure fell below the systolic minimum setting | O
MDC_DEV_ANALY_SAT_O2_VMD
MDC_DEV_ANALY_SAT_O2_CHAN
MDC_PULS_OXIM_SAT_O2 | MDC_EVT_HI | phys high thr | Oximeter Oxygen Saturation High | Notification that Oxygen Saturation is above limit. | O
MDC_PULS_OXIM_SAT_O2 | MDC_EVT_LO | phys low thr | Oximeter Oxygen Saturation Low | Notification that Oxygen Saturation is below limit.
MDC_PULS_OXIM_PULS_RATE | MDC_EVT_HI | phys high thr | Oximeter Pulse High | Notification that the pulse rate measure by the pulse oximeter is above limit. | O
MDC_PULS_OXIM_PULS_RATE | MDC_EVT_LO | phys low thr | Oximeter Pulse Low | Notification that the pulse rate measure by the pulse oximeter is below limit.
MDC_DEV_ANALY_SAT_O2_CHAN | MDCX_EVT_PULS_OXIM_ERROR | tech | Pulse Oximeter Error | Notification of a detected error on the pulse oximeter device | O
MDCX_DEV_BLOOD_CHEM_VMD
MDCX_DEV_BLOOD_CHEM_CHAN
MDC_CONC_HCT | MDC_EVT_HI | phys high thr | Hematocrit High | Notification that patient hematocrit has risen above limit. | O
MDC_CONC_HCT | MDC_EVT_LO | phys low thr | Hematocrit Low | Notification that patent hematocrit has fallen below limit. | O
MDC_CONC_HB | MDC_EVT_HI | phys high thr | Hemoglobin High | Notification that patient hemoglobin has risen above limit. | O
MDC_CONC_HB | MDC_EVT_LO | phys low thr | Hemoglobin Low | Notification that patent hemoglobin has fallen below limit. | O
MDCX_HDIALY_REL_BLOOD_VOL | MDC_EVT_LO | phys low thr | Blood Volume | Notification that the relative blood volume has dropped below the defined setting | O
MDC_SAT_O2 | MDC_EVT_HI | phys high thr | Oxygen Saturation High | Notification that Patient Oxygen Saturation has risen above limit. | O
MDC_SAT_O2 | MDC_EVT_LO | phys low thr | Oxygen Saturation Low | Notification that Patient Oxygen Saturation has fallen below limit. | O

## ConditionCodes
Code | Condition | Expression
C1 | If Anticoagulation Mode is not None | MDC_HDIALY_ANTICOAG_MODE != NONE
C2 | If Anticoagulation Mode includes Bolus | MDC_HDIALY_ANTICOAG_MODE = BOL
C3 | If Anticoagulation Mode includes Continuous or Profile | MDC_HDIALY_ANTICOAG_MODE = ( CON, PRO )
C4 | One for each anticoagulant pump | N/A
C5 | If Blood Pump Mode is not double needle | MDC_HDIALY_BLD_PUMP_MODE != 2N
C6 | If Mode is HD or HDF | MDC_HDIALY_MACH_MODE_OF_OPERATION = ( HD, HDF )
C7 | If Mode is HF or HDF and Dilution is Post or Pre-Post | MDC_HDIALY_MACH_MODE_OF_OPERATION = ( HD, HDF ) AND MDC_HDIALY_RF_DILUTION_LOCATION = ( POSTF, PREF-POSTF )
C8 | If Mode is HDF or HF | MDC_HDIALY_MACH_MODE_OF_OPERATION = ( HD, HF )
C9 | If Replacement Fluid Mode is Bolus | MDC_HDIALY_RF_FLOW_MODE = BOLUS
C10 | If Completion Method is not None | MDC_HDIALY_THERAPY_COMPLETE_METHOD != NONE
C11 | if UF removal includes a target volume | exists (MDC_HDIALY_UF_TARGET_VOL_TO_REMOVE)
C12 | If Mode is HD or HDF and the dialysis machine knows the conductivity through direct measurement or by knowing what type of bagged fluid is being used. | MDC_HDIALY_MACH_MODE_OF_OPERATION = ( HDF ) AND ( exists (MDC_HDIALY_DIALYSATE_COND ) OR ( exists (bagged fluid conductivity) ) [not finalized at present time]
C13 | If alarm limit is being set.
C14 | If Completion Method (MDC_HDIALY_THERAPY_COMPLETE_METHOD) is Clock Time. | MDC_HDIALY_THERAPY_COMPLETE_METHOD = CT
C15 | If Completion Method (MDC_HDIALY_THERAPY_COMPLETE_METHOD) is Treatment Time. | MDC_HDIALY_THERAPY_COMPLETE_METHOD = AT
C16 | If Mode is HF or HDF and Dilution is Pre or Pre-Post | MDC_HDIALY_MACH_MODE_OF_OPERATION = ( HD, HDF ) AND MDC_HDIALY_RF_DILUTION_LOCATION = ( PREF, PREF-POSTF )
C17 | If Completion Method (MDC_HDIALY_THERAPY_COMPLETE_METHOD) is Kt/V. | MDC_HDIALY_THERAPY_COMPLETE_METHOD = KTV
C18 | If UF Target (MDC_HDIALY_UF_TARGET_VOL_TO_REMOVE) or Treatment Time (MDC_HDIALY_THERAPY_TX_TIME_SETTING) is not present.
C19 | If UF Mode (MDC_HDIALY_UF_MODE) is profiled. | MDC_HDIALY_UF_MODE = (PRO-WT, PRO-WOT)
C20 | If Anticoagulant Mode (MDC_HDIALY_ANTICOAG_MODE) is profiled. | MDC_HDIALY_ANTICOAG_MODE = (PRO, BOLPRO)
C21 | If Blood Pump Mode (MDC_HDIALY_BLD_PUMP_MODE) is profiled. | TBD
C22 | If Dialysate Mode (MDC_HDIALY_DIALYSATE_FLOW_MODE) is profiled. | MDC_HDIALY_DIALYSATE_FLOW_MODE = PRO
C23 | If Sodium Mode (MDC_HDIALY_DIALYSATE_CONC_NA_MODE) is profiled. | MDC_HDIALY_DIALYSATE_CONC_NA_MODE = PRO
C24 | If RF Delivery Mode (MDC_HDIALY_RF_FLOW_MODE) is profiled. | MDC_HDIALY_RF_FLOW_MODE = PRO