## Data
REFID | SysName | SysNameBase | SysName1stDesc | SysName2ndDesc | SysName3rdDesc | AlertType | CommonTerm | Description | Phase | Temporal | DataType | Format | UOM | UOM_UCUM | Use | RxUse | Enum_Table | Enum_Values | Enum_Descriptions | Range | Example | gp | s | Disc | PART | CODE10 | PartCode | CF_CODE10
MDC_DEV_PDIALY_MACHINE_MDS | System ¦ ¦ ¦ Dialysis, Peritoneal | System |  |  | Dialysis, Peritoneal |  | Peritoneal Dialysis machine | A device or system used to treat renal failure delivered by removing toxins from the blood through an artificial kidney, called a dialyzer. | All | All |  |  |  |  | M | M | [5 empty] | 1 |  | MVC(1) | 1 | 5473 | 1:5473 | 71009
. . . MDC_ID_MODEL_MANUFACTURER | System (string) ¦ Identifier ¦ Manufacturer ¦ Device | System (string) | Identifier | Manufacturer | Device |  | Dialysis Device Manufacturer | String containing device manufacturer's name | All | All | String | Alphanumeric | N/A | -- | M | X | -- |  | -- | N/A | NxStage | 0 |  | 1(0) | 8 | 7682 | 8:7682 | 531970
. . . MDC_ID_MODEL_NUMBER | System (string) ¦ Identifier ¦ Model Number ¦ Device | System (string) | Identifier | Model Number | Device |  | Dialysis Device Model | String containing device manufacturer's model identifier for the device | All | All | String | Alphanumeric | N/A | -- | M | X | -- |  | -- | N/A | System One | 0 |  | 1(0) | 8 | 7681 | 8:7681 | 531969
. . . MDC_ID_PROD_SPEC_SERIAL | System (string) ¦ Identifier ¦ Serial Number ¦ Device | System (string) | Identifier | Serial Number | Device |  | Dialysis Device Serial Number | String containing the device’s serial number | All | All | String | Alphanumeric | N/A | -- | M | X | -- |  | -- | N/A | AS1000478 | 0 |  | 1(0) | 8 | 7684 | 8:7684 | 531972
. . . MDC_ID_PROD_SPEC_SW | System (string) ¦ Identifier ¦ Software Version ¦ Device | System (string) | Identifier | Software Version | Device |  | Device Software Version | String defining the version of the software on the machine | All | All | String | Alphanumeric | N/A | -- | M | X | -- |  | -- | N/A | 1.2.3.4 | 0 |  | 1(0) | 8 | 7687 | 8:7687 | 531975
. . . MDC_ATTR_ID_UDI | System (string) ¦ Identifier, Unique ¦ Universal identifier, UDI ¦ Device | System (string) | Identifier, Unique | Universal identifier, UDI | Device |  | Dialysis Device UDI | String containing the device’s FDA required Universal Device Identifier . [Note: MDC_ATTR_ID_UDI contains three components: UdiAuthority, UdiIssuer, UdiLabel to support non-FDA entities, but currently there is no approved way of encoding this using HL7 V2; PRT-10 and PRT-16-20 is used instead; UdiLabel = PRT-10.] | All | All | String | Alphanumeric | N/A | -- | O | X | -- |  | -- | N/A | +M535NX10003A0/$$+735241/16D20180305J | 0 |  | 1(0) | 1 | 2380 | 1:2380 | 67916
. MDC_DEV_PDIALY_VMD | Device ¦ ¦ ¦ Dialysis, Peritoneal | Device |  |  | Dialysis, Peritoneal |  | Peritoneal Dialysis subsystem | A device or subsystem used to treat renal failure delivered by removing toxins from the blood through an artificial kidney, called a dialyzer. | All | All |  |  |  |  | M | M | [5 empty] | 1 |  | MVC(2) | 1 | 5478 | 1:5478 | 71014
. . MDC_DEV_PDIALY_MACH_CONFIG_CHAN | Subsystem ¦ ¦ Machine Configuration ¦ Dialysis, Peritoneal | Subsystem |  | Machine Configuration | Dialysis, Peritoneal |  | Peritoneal Dialysis configuration | The principal dialysis machine operational modes, treatment modality, settings and other information about the dialysis machine. | All | All |  |  |  |  | M | X | [5 empty] | 1 |  | MVC(3) | 1 | 5483 | 1:5483 | 71019
. . . MDC_ATTR_TIME_ABS | Time ¦ Machine Time ¦ Machine Configuration ¦ Dialysis | Time | Machine Time | Machine Configuration | Dialysis |  | Dialysis Device Time | Date and time as recorded on the dialysis device's internal clock with offset to UTC (based on location and DST) | All | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | M | X | -- |  | -- | N/A | 20190421130045+/-0000 | 0 |  | 1(0) | 1 | 2439 | 1:2439 | 67975
. . MDC_DEV_PDIALY_TREATMENT_CHAN | Subsystem ¦ Treatment ¦ ¦ Dialysis, Peritoneal | Subsystem | Treatment |  | Dialysis, Peritoneal |  | Treatment | Treatment channel | Intradialytic | All |  |  |  |  | M | M | [5 empty] | 1 |  | MVC(3) | 1 | 5487 | 1:5487 | 71023
. . . MDC_PDIALY_TREAT_TYPE | Treatment Type ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Treatment Type | Setting | Treatment | Dialysis, Peritoneal |  | Treatment Type | Treatment Type | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | M | M | PD_TBL_01 | CAPD APD CCPD NIPD PDplus TPD aAPD IPD Other | CAPD - Continuous Ambulatory Peritoneal Dialysis APD _ Automated Peritoneal Dialysis CCPD - Continuous Cycling Peritoneal Dialysis NIPD - Nocturnal Intermittent Peritoneal Dialysis PDPLUS - PD Plus TPD - Tidal Peritoneal Dialysis AAPD - Adapted Automated Peritoneal Dialysis IPD - Intermittent Peritoneal Dialysis OTHER - Undefined | N/A | CAPD | 1 |  | 1(0) | 2 | 27561 | 2:27561 | 158633
. . . MDC_PDIALY_TREAT_LOCATION | Treatment Location ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | Treatment Location | Status | Treatment | Dialysis, Peritoneal |  | Treatment Location | Treatment Location | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | O | O | PD_TBL_02 | HOME CLINIC OTHER | Home Clinic Other/Unknown | N/A | HOME | 1 |  | 1(0) | 2 | 27562 | 2:27562 | 158634
. . . MDC_PDIALY_CURRENT_PHASE | Current Phase ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | Current Phase | Status | Treatment | Dialysis, Peritoneal |  | Current Phase | If this value appears in an exchange channel, it is the phase of the exchange. If the value appears in the treatment channel, it is the phase of the current exchange. | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | M | X | PD_TBL_03 | PENDING FILL DWELL DRAIN COMPLETE | Pending Fill Dwell Drain Complete | N/A | COMPLETE | 1 |  | 1(0) | 2 | 27563 | 2:27563 | 158635
. . . MDC_PDIALY_PRESCRIPTION_TYPE | Type ¦ Status ¦ Prescription ¦ Dialysis, Peritoneal | Type | Status | Prescription | Dialysis, Peritoneal |  | Prescription Type | Prescription Type | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | M | M | PD_TBL_13 | EXCH TXDUR TXEND | Exchange Based Treatment Based - Fixed Duration Treatment Based - Fixed End Time | N/A | EXCH | 1 |  | 1(0) | 2 | 27564 | 2:27564 | 158636
. . . MDC_PDIALY_APD_START_TIME_SETTING | ADP Start Time ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | ADP Start Time | Status | Treatment | Dialysis, Peritoneal |  | APD Start Time | APD Start Time | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | C5 | C11 | N/A |  |  |  | 20190421130045+/-0000 | 1 |  | 1(0) | 258 | 27565 | 258:27565 | 16935853
. . . MDC_PDIALY_APD_END_TIME_SETTING | ADP End Time ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | ADP End Time | Status | Treatment | Dialysis, Peritoneal |  | APD End Time | APD End Time | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | C5 | C12 | N/A |  |  |  | 20190421130045+/-0000 | 1 |  | 1(0) | 258 | 27566 | 258:27566 | 16935854
. . . MDC_PDIALY_APD_TOTAL_VOLUME_SETTING | ADP Volume ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | ADP Volume | Status | Treatment | Dialysis, Peritoneal |  | APD Total Volume | Total volume in APD exchanges. | Intradialytic | All | Numeric | X.XXX | L | L | C5 | C12 | N/A |  |  |  |  | 1 |  | 1(0) | 258 | 27567 | 258:27567 | 16935855
. . . MDC_PDIALY_APD_CYCLES_SETTING | ADP Cycler ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | ADP Cycler | Setting | Treatment | Dialysis, Peritoneal |  | APD Cycles | APD Cycles | Intradialytic | All | Numeric | XX | N/A | -- | C5 | C12 | N/A |  |  |  |  | 1 |  | 1(0) | 258 | 27568 | 258:27568 | 16935856
. . . MDC_PDIALY_CAPD_START_TIME_SETTING | CAPD Start Time ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | CAPD Start Time | Status | Treatment | Dialysis, Peritoneal |  | CAPD Start Time | CAPD Start Time | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | C6 | C13 | N/A |  |  |  | 20190421130045+/-0000 | 1 |  | 1(0) | 258 | 27569 | 258:27569 | 16935857
. . . MDC_PDIALY_CAPD_END_TIME_SETTING | CAPD End Time ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | CAPD End Time | Status | Treatment | Dialysis, Peritoneal |  | CAPD End Time | CAPD End Time | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | C6 | C14 | N/A |  |  |  | 20190421130045+/-0000 | 1 |  | 1(0) | 258 | 27570 | 258:27570 | 16935858
. . . MDC_PDIALY_CAPD_TOTAL_VOLUME_SETTING | CAPD Volume ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | CAPD Volume | Status | Treatment | Dialysis, Peritoneal |  | CAPD Exchanges | CAPD Exchanges | Intradialytic | All | Numeric |  | N/A | -- | C6 | C14 | N/A |  |  |  |  | 1 |  | 1(0) | 258 | 27571 | 258:27571 | 16935859
. . . MDC_PDIALY_CAPD_CYCLES_SETTING | CAPD Cycles ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | CAPD Cycles | Setting | Treatment | Dialysis, Peritoneal |  | CAPD Total Volume | Total volume in CAPD exchanges | Intradialytic | All | Numeric | XX | N/A | -- | C6 | C14 | N/A |  |  |  |  | 1 |  | 1(0) | 258 | 27572 | 258:27572 | 16935860
. . . MDC_PDIALY_TIDAL_PERCENTAGE | Tidal Percentage ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Tidal Percentage | Setting | Treatment | Dialysis, Peritoneal |  | Tidal Percentage | The percentage of Night Fill Volume to be drained during a tidal cycle. | Intradialytic | All | Numeric | XXX.X | % | % | O | O | N/A |  |  |  |  | 1 |  | 1(0) | 2 | 27573 | 2:27573 | 158645
. . . MDC_PDIALY_FULL_DRAIN_FREQ | Full Drain Frequency ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Full Drain Frequency | Setting | Treatment | Dialysis, Peritoneal |  | Full Drain Frequency | The number of cycles between full drains when performing tidal exchanges. The values is expressed as the number of drains between full drains. | Intradialytic | All | Numeric | X | N/A | -- | O | O | N/A |  |  |  |  | 1 |  | 1(0) | 2 | 27574 | 2:27574 | 158646
. . . MDC_PDIALY_PERITONEUM_VOL_MAX | Peritoneum Max Volume ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Peritoneum Max Volume | Setting | Treatment | Dialysis, Peritoneal |  | Maximum Peritoneum Volume | The maximum prescribed amount of dialysate to fill the peritoneal cavity. | Intradialytic | All | Numeric | X.XXX | L | L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27575 | 2:27575 | 158647
. . . MDC_PDIALY_INIT_DRAIN_VOL_MIN_SETTING | Drain Min Volume ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Drain Min Volume | Setting | Treatment | Dialysis, Peritoneal |  | Minimum Initial Drain Volume | The minimum volume to remove during the initial drain. | Intradialytic | All | Numeric | X.XXX | L | L | O | O | N/A |  | N/A |  |  | 1 |  | MMM(2) | 258 | 28162 | 258:28162 | 16936450
. . . MDC_PDIALY_INIT_DRAIN_VOL_MAX_SETTING | Drain Max Volume ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Drain Max Volume | Setting | Treatment | Dialysis, Peritoneal |  | Maximum Initial Drain Time | The maximum volume to remove during the initial drain. | Intradialytic | All | Numeric | X.XXX | L | L | O | O | N/A |  | N/A |  |  | 1 |  | MMM(1) | 258 | 28161 | 258:28161 | 16936449
. . . MDC_PDIALY_SAMPLE_REMINDER | Sample Reminder Required ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Sample Reminder Required | Setting | Treatment | Dialysis, Peritoneal |  | Sample Reminder | This value is true when the user should be reminded to collect an effluent sample. | Intradialytic | All | Bool | T / F | N/A | -- | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27576 | 2:27576 | 158648
. . . MDC_PDIALY_MEDICATION_SURVEY | Medication Survey Required ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Medication Survey Required | Setting | Treatment | Dialysis, Peritoneal |  | Medications Survey | This value is true when the user should be surveyed for medications taken. | Intradialytic | All | Bool | T / F | N/A | -- | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27577 | 2:27577 | 158649
. . . MDC_PDIALY_EXCHANGE_SURVEY | Exchange Survey Required ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Exchange Survey Required | Setting | Treatment | Dialysis, Peritoneal |  | Exchange Survey | This value is true when the user should be surveyed for number of day exchanges performed. | Intradialytic | All | Bool | T / F | N/A | -- | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27578 | 2:27578 | 158650
. . . MDC_PDIALY_DAY_DRAIN_TIME_MIN_SETTING | Minimum Drain Time ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Minimum Drain Time | Setting | Treatment | Dialysis, Peritoneal |  | Minimum Day Drain Time | Minimum duration for a day time drain. | Intradialytic | All | Numeric | XXX.X | min | min | O | O | N/A |  | N/A |  |  | 1 |  | MMM(2) | 258 | 28166 | 258:28166 | 16936454
. . . MDC_PDIALY_DAY_DRAIN_VOL_PCT_MIN_SETTING | Minimum Drain Volume Percentage ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Minimum Drain Volume Percentage | Setting | Treatment | Dialysis, Peritoneal |  | Minimum Day Drain Volume Percent | The min amount of fluid to drain expressed as a percentage of the fill volume. | Intradialytic | All | Numeric | XXX.X | % |  | O | O | N/A |  | N/A |  |  | 1 |  | MMM(2) | 258 | 28170 | 258:28170 | 16936458
. . . MDC_PDIALY_NIGHT_DRAIN_TIME_MIN_SETTING | Minimum Night Drain Time ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Minimum Night Drain Time | Setting | Treatment | Dialysis, Peritoneal |  | Minimum Night Drain Time | Minimum duration for a night time drain. | Intradialytic | All | Numeric | XXX.X | min | min | O | O | N/A |  | N/A |  |  | 1 |  | MMM(2) | 258 | 28174 | 258:28174 | 16936462
. . . MDC_PDIALY_LAST_DRAIN_MODE | Last Drain Mode ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Last Drain Mode | Setting | Treatment | Dialysis, Peritoneal |  | Last Drain Mode | This value specifies how the last drain phase is performed. | Intradialytic | All | Enumeration |  | N/A | -- | O | O | PD_TBL_12 | NEVER ALWAYS UFLIMIT | Never - It will never take place Always - It will always take place UF Limit - It will take place if the total estimated UF for the therapy is below the Extra Drain UF Limit. The Extra Drain UF Limit is expressed as a percentage of total estimated UF for the therapy. |  |  | 1 |  | 1(0) | 2 | 27579 | 2:27579 | 158651
. . . MDC_PDIALY_LAST_DRAIN_UF_LIMIT | Last Drain UF Limit ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Last Drain UF Limit | Setting | Treatment | Dialysis, Peritoneal |  | Last Drain UF Limit | This limit determines if the last drain is performed. The Last Drain UF Limit is expressed as a percentage of total estimated UF divided by the total drain volume. If the value is below the Last Drain UF Limit the last drain is performed. | Intradialytic | All | \ | XXX.X | % | % | O | O | N/A |  |  |  |  | 1 |  | 1(0) | 2 | 27580 | 2:27580 | 158652
. . . MDC_PDIALY_NIGHT_UF_ESTIMATE | Night UF Estimate ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | Night UF Estimate | Status | Treatment | Dialysis, Peritoneal |  | Estimated Night UF | The estimated UF volume for the night therapy. This value is used to keep track of the fluid balance for the patient, so that the fluid that was filled plus the fluid generated as UF is drained from the patient. | Intradialytic | All | Numeric | X.XXX | L | L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27581 | 2:27581 | 158653
. . . MDC_PDIALY_TOTAL_UF_ESTIMATE | [6 empty] | Total UF | The estimated UF volume drained from the patient during the course of the treatment. Total UF is the sum of the UF for all of the indivial exchanges. The Initial Drain and Last Fill Volume are not included in the Total UF. |  |  | Numeric | X.XXX | L | L | O | O | N/A | [6 empty] | 1(0) | 2 | 27582 | 2:27582 | 158654
. . . MDC_PDIALY_TEMP_DIALYSATE_SETTING | Dialysate Temperature ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Dialysate Temperature | Setting | Treatment | Dialysis, Peritoneal |  | Dialysate Temperature Setting | Desired dialysate temperature during treatment | Intradialytic | All | Numeric | XX.X | ˚C | Cel | O | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 258 | 27583 | 258:27583 | 16935871
. . MDC_DEV_PDIALY_FLUID_1_CHAN . . MDC_DEV_PDIALY_FLUID_2_CHAN . . MDC_DEV_PDIALY_FLUID_3_CHAN . . MDC_DEV_PDIALY_FLUID_4_CHAN . . MDC_DEV_PDIALY_FLUID_5_CHAN . . MDC_DEV_PDIALY_FLUID_6_CHAN | Subsystem ¦ Fluid ¦ ¦ Dialysis, Peritoneal | Subsystem | Fluid |  | Dialysis, Peritoneal |  | Fluid source line or bag | Fluid source line or bag channel | Intradialytic | All: All; O: M; O: M |  |  |  |  | 1 1 1 1 1 1 |  | MVC(3) MVC(3) MVC(3) MVC(3) MVC(3) MVC(3) | 1 1 1 1 1 1 | 5491 5495 5499 5503 5507 5511 | 1::5491 1::5495 1::5499 1::5503 1::5507 1::5511 | 71027 71031 71035 71039 71043 71047
. . . MDC_PDIALY_FLUID_NAME | Name ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Name | Setting | Fluid | Dialysis, Peritoneal |  | Dialysate Name | Dialysate Name | Intradialytic | All | String | Alphanumeric | N/A | -- | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27584 | 2:27584 | 158656
. . . MDC_DIALY_FLUID_GLUCOSE_CONC | Glucose ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Glucose | Setting | Fluid | Dialysis, Peritoneal |  | Glucose | Glucose | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27585 | 2:27585 | 158657
. . . MDC_DIALY_FLUID_ICODEXTRIM_CONC | Icodextrim ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Icodextrim | Setting | Fluid | Dialysis, Peritoneal |  | Icodextrin | Icodextrin | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27586 | 2:27586 | 158658
. . . MDC_DIALY_FLUID_AMINO_ACID_CONC | Amino Acid ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Amino Acid | Setting | Fluid | Dialysis, Peritoneal |  | Amino acid | Amino acid | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27587 | 2:27587 | 158659
. . . MDC_DIALY_FLUID_OSMOTIC_AGENT | Agent ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Agent | Setting | Fluid | Dialysis, Peritoneal |  | Other osmotic agent type | Other osmotic agent type | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | O | O | PD_TBL_04 | GLU DEX ICO OTHER | Glucose Dextrose Icodextrin Other |  |  | 1 |  | 1(0) | 2 | 27588 | 2:27588 | 158660
. . . MDC_DIALY_FLUID_OSMOTIC_AGENT_CONC | Agent Concentration ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Agent Concentration | Setting | Fluid | Dialysis, Peritoneal |  | Other osmotic agent concentration | Other osmotic agent concentration | Intradialytic | All | Numeric | XX.XX | % |  | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27589 | 2:27589 | 158661
. . . MDC_DIALY_FLUID_BICARB_CONC | Bicarb ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Bicarb | Setting | Fluid | Dialysis, Peritoneal |  | Bicarbonate (HCO3) | Bicarbonate (HCO3) | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27590 | 2:27590 | 158662
. . . MDC_DIALY_FLUID_CHLORIDE_CONC | Chloride ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Chloride | Setting | Fluid | Dialysis, Peritoneal |  | Chloride (CL) | Chloride (CL) | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27591 | 2:27591 | 158663
. . . MDC_DIALY_FLUID_LACTATE_CONC | Lacatate ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Lacatate | Setting | Fluid | Dialysis, Peritoneal |  | Lactate | Lactate | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27592 | 2:27592 | 158664
. . . MDC_DIALY_FLUID_CALCIUM_CONC | Calcium ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Calcium | Setting | Fluid | Dialysis, Peritoneal |  | Calcium (CA) | Calcium (CA) | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27593 | 2:27593 | 158665
. . . MDC_DIALY_FLUID_CITRATE_CONC | Citrate ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Citrate | Setting | Fluid | Dialysis, Peritoneal |  | Citrate | Citrate | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27594 | 2:27594 | 158666
. . . MDC_DIALY_FLUID_POTASSIUM_CONC | Potassium ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Potassium | Setting | Fluid | Dialysis, Peritoneal |  | Potassium (K) | Potassium (K) | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27595 | 2:27595 | 158667
. . . MDC_DIALY_FLUID_MAGNESIUM_CONC | Magnesium ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Magnesium | Setting | Fluid | Dialysis, Peritoneal |  | Magnesium | Magnesium | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27596 | 2:27596 | 158668
. . . MDC_PDIALY_FLUID_BAG_VOLUME | Volume ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Volume | Setting | Fluid | Dialysis, Peritoneal |  | Bag Volume | Bag Volume | Intradialytic | All | Numeric | X.XXX | L | L | C4 | C4 | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27597 | 2:27597 | 158669
. . . MDC_PDIALY_FLUID_SOURCE | Source ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Source | Setting | Fluid | Dialysis, Peritoneal |  | Fluid Source | Source of the fluid. | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | M | M | PD_TBL_05 | BAG ONLINE UNKNOWN | Bag Online Uknown |  |  | 1 |  | 1(0) | 2 | 27598 | 2:27598 | 158670
. . MDC_DEV_PDIALY_EXCHANGE_CHAN | Subsystem ¦ Exchange ¦ ¦ Dialysis, Peritoneal | Subsystem | Exchange |  | Dialysis, Peritoneal |  | Peritoneal Exchange | The values for the current and previous exchanges. | Intradialytic | All: All; O: M; O: C7 |  |  |  |  | 1 |  | MVC(3) | 1 | 5515 | 1:5515 | 71051
. . . MDC_ATTR_CHAN_NUM_LOGICAL | Number ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Number | Setting | Exchange | Dialysis, Peritoneal |  | Exchange Number | The number of the exchange in the current treatment. The first exchange is number 1. | Intradialytic | All | Numeric | X | N/A | -- | M | C7 | -- |  | -- | 1-9 | 1 | 0 |  | 1(0) | 1 | 2606 | 1:2606 | 68142
. . . MDC_PDIALY_CURRENT_PHASE | Phase ¦ Status ¦ Exchange ¦ Dialysis, Peritoneal | Phase | Status | Exchange | Dialysis, Peritoneal |  | Current Phase | If this value appears in an exchange channel, it is the phase of the exchange. If the value appears in the treatment channel, it is the phase of the current exchange. | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | M | X | PD_TBL_03 |  | See MDC_PDIALY_CURRENT_PHASE value in the Treatment Channel. |  |  | 1 |  | 1(0) | 2 | 27563 | 2:27563 | 158635
. . . MDC_PDIALY_FILL_START_TIME | Fill Start ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Fill Start | Time | Exchange | Dialysis, Peritoneal |  | Fill Start time | The time that the fill phase started. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | C1 | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27599 | 2:27599 | 158671
. . . MDC_PDIALY_FILL_END_TIME | Fill End ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Fill End | Time | Exchange | Dialysis, Peritoneal |  | Fill End time | The time that the fill phase ended. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | C1 | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27600 | 2:27600 | 158672
. . . MDC_PDIALY_FILL_MODE | Fill Mode ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Fill Mode | Setting | Exchange | Dialysis, Peritoneal |  | Fill mode | Fill mode | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | C1 | C7 | PD_TBL_06 | AUTO MANUAL | Cycler Manual |  |  | 1 |  | 1(0) | 2 | 27601 | 2:27601 | 158673
. . . MDC_PDIALY_FILL_DURATION_SETTING | Fill Duration ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Fill Duration | Setting | Exchange | Dialysis, Peritoneal |  | Prescribed Fill Duration | Prescribed Fill duration | Intradialytic | All | Numeric | XXX.X | min | min | C1 | C7 | N/A |  | N/A |  |  | 1 |  | 1(0) | 258 | 27602 | 258:27602 | 16935890
. . . MDC_PDIALY_FILL_DURATION | Fill Duration ¦ Actual ¦ Exchange ¦ Dialysis, Peritoneal | Fill Duration | Actual | Exchange | Dialysis, Peritoneal |  | Actual Fill Duration | Actual fill duration | Intradialytic | All | Numeric | XXX.X | min | min | C1 | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27602 | 2:27602 | 158674
. . . MDC_PDIALY_FILL_VOLUME_SETTING | Fill Volume ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Fill Volume | Setting | Exchange | Dialysis, Peritoneal |  | Prescribed Fill volume | Prescribed Fill volume | Intradialytic | All | Numeric Array | X.XXX | L | L | C1 | C7 | N/A |  | N/A |  |  | 1 |  | 1(0) | 258 | 27603 | 258:27603 | 16935891
. . . MDC_PDIALY_FILL_VOLUME | Fill Volume ¦ Actual ¦ Exchange ¦ Dialysis, Peritoneal | Fill Volume | Actual | Exchange | Dialysis, Peritoneal |  | Actual Fill volume | Actual Fill volume | Intradialytic | All | Numeric Array | X.XXX | L | L | C1 | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27603 | 2:27603 | 158675
. . . MDC_PDIALY_FLUID_TEMP_SETTING | Fluid Temperature ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Fluid Temperature | Setting | Exchange | Dialysis, Peritoneal |  | Fluid temperature setting | Fluid temperature | Intradialytic | All | Numeric | XX.X | degC | Cel | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 258 | 27604 | 258:27604 | 16935892
. . . MDC_PDIALY_FLUID_TEMP | Fluid Temperature ¦ Status ¦ Exchange ¦ Dialysis, Peritoneal | Fluid Temperature | Status | Exchange | Dialysis, Peritoneal |  | Fluid temperature | Fluid temperature | Intradialytic | All | Numeric | XX.X | degC | Cel | O | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27604 | 2:27604 | 158676
. . . MDC_PDIALY_DWELL_START_TIME | Dwell Start ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Dwell Start | Time | Exchange | Dialysis, Peritoneal |  | Dwell Start time | The time that the dwell phase started. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | C2 | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27605 | 2:27605 | 158677
. . . MDC_PDIALY_DWELL_END_TIME | Dwell End ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Dwell End | Time | Exchange | Dialysis, Peritoneal |  | Dwell End time | The time that the dwell phase ended. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | C2 | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27606 | 2:27606 | 158678
. . . MDC_PDIALY_DWELL_DURATION_SETTING | Dwell Duration ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Dwell Duration | Setting | Exchange | Dialysis, Peritoneal |  | Prescribed Dwell Duration | Prescribed dwell duration | Intradialytic | All | Numeric | XXX.X | min | min | M | C7 | N/A |  | N/A |  |  | 1 |  | 1(0) | 258 | 27607 | 258:27607 | 16935895
. . . MDC_PDIALY_DWELL_DURATION | Dwell Duration ¦ Actual ¦ Exchange ¦ Dialysis, Peritoneal | Dwell Duration | Actual | Exchange | Dialysis, Peritoneal |  | Actual Dwell Duration | Actual dwell duration | Intradialytic | All | Numeric | XXX.X | min | min | M | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27607 | 2:27607 | 158679
. . . MDC_PDIALY_DRAIN_START_TIME | Drain Start ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Drain Start | Time | Exchange | Dialysis, Peritoneal |  | Drain Start time | The time that the drain phase started. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | C3 | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27608 | 2:27608 | 158680
. . . MDC_PDIALY_DRAIN_END_TIME | Drain End ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Drain End | Time | Exchange | Dialysis, Peritoneal |  | Drain End time | The time that the drain phase ended. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | C3 | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27609 | 2:27609 | 158681
. . . MDC_PDIALY_DRAIN_MODE | Drain Mode ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Drain Mode | Setting | Exchange | Dialysis, Peritoneal |  | Drain mode | Drain mode | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | C3 | C7 | PD_TBL_07 | TOBAG TODRAIN UNKNOWN | Cycler to Bag Cycler to drain Manual |  |  | 1 |  | 1(0) | 2 | 27610 | 2:27610 | 158682
. . . MDC_PDIALY_DRAIN_DURATION_SETTING | Drain Duration ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Drain Duration | Setting | Exchange | Dialysis, Peritoneal |  | Prescribed Drain duration | Prescribed drain duration | Intradialytic | All | Numeric | XXX.X | min | min | C3 | C7 | N/A |  | N/A |  |  | 1 |  | 1(0) | 258 | 27611 | 258:27611 | 16935899
. . . MDC_PDIALY_DRAIN_DURATION | Drain Duration ¦ Actual ¦ Exchange ¦ Dialysis, Peritoneal | Drain Duration | Actual | Exchange | Dialysis, Peritoneal |  | Actual Drain duration | Actual drain duration | Intradialytic | All | Numeric | XXX.X | min | min | C3 | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27611 | 2:27611 | 158683
. . . MDC_PDIALY_DRAIN_VOLUME_SETTING | Drain Volume ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Drain Volume | Setting | Exchange | Dialysis, Peritoneal |  | Prescribed Drain volume | Prescribed Drain volume | Intradialytic | All | Numeric | X.XXX | L | L | C3 | C7 | N/A |  | N/A |  |  | 1 |  | 1(0) | 258 | 27612 | 258:27612 | 16935900
. . . MDC_PDIALY_DRAIN_VOLUME | Drain Volume ¦ Actual ¦ Exchange ¦ Dialysis, Peritoneal | Drain Volume | Actual | Exchange | Dialysis, Peritoneal |  | Actual Drain volume | Actual Drain volume | Intradialytic | All | Numeric | X.XXX | L | L | C3 | X | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27612 | 2:27612 | 158684
. . MDC_DEV_PDIALY_SUPPLY_LINE_CHAN | Subsystem ¦ Supply Line ¦ ¦ Dialysis, Peritoneal | Subsystem | Supply Line |  | Dialysis, Peritoneal |  | Fluid supply line | Fluid supply line channel | All | All: All; O: O; O: X |  |  |  |  | 1 |  | MVC(3) | 1 | 5519 | 1:5519 | 71055
. . . MDC_PDIALY_PRESSURE_DIALYSATE | Pressure ¦ Status ¦ Supply Line ¦ Dialysis, Peritoneal | Pressure | Status | Supply Line | Dialysis, Peritoneal |  | Dialysate Supply Pressure Sensor Value | Dialysate pressure in the supply line(s). | Intradialytic | All | Numeric | XXXX | mmHg | mm[Hg] | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28176 | 2:28176 | 159248
. . MDC_DEV_PDIALY_DRAIN_LINE_CHAN | Subsystem ¦ Drain Line ¦ ¦ Dialysis, Peritoneal | Subsystem | Drain Line |  | Dialysis, Peritoneal |  | Fluid drain line | Fluid drain line channel | All | All: All; O: O; O: X |  |  |  |  | 1 |  | MVC(3) | 1 | 5523 | 1:5523 | 71059
. . . MDC_PDIALY_PRESSURE_DRAIN | Pressure ¦ Status ¦ Drain Line ¦ Dialysis, Peritoneal | Pressure | Status | Drain Line | Dialysis, Peritoneal |  | Drain Pressure Sensor Value | Dialysate pressure in the drain line. | Intradialytic | All | Numeric | XXXX | mmHg | mm[Hg] | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28180 | 2:28180 | 159252
. . MDC_DEV_PDIALY_PATIENT_LINE_CHAN | Subsystem ¦ Patient Line ¦ ¦ Dialysis, Peritoneal | Subsystem | Patient Line |  | Dialysis, Peritoneal |  | Fluid patient line | Fluid patient line channel | All | All: All; O: O; O: X |  |  |  |  | 1 |  | MVC(3) | 1 | 5527 | 1:5527 | 71063
. . . MDC_PDIALY_PRESSURE_PATIENT | Pressure ¦ Status ¦ Patient Line ¦ Dialysis, Peritoneal | Pressure | Status | Patient Line | Dialysis, Peritoneal |  | Patient Pressure Sensor Value | Dialysate pressure in the patient line. | Intradialytic | All | Numeric | XXXX | mmHg | mm[Hg] | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28184 | 2:28184 | 159256
. . MDC_DEV_PDIALY_HEATER_LINE_CHAN | Subsystem ¦ Heater Line ¦ ¦ Dialysis, Peritoneal | Subsystem | Heater Line |  | Dialysis, Peritoneal |  | Fluid heater line | Fluid heater line channel | All | All: All; O: O; O: X |  |  |  |  | 1 |  | MVC(3) | 1 | 5531 | 1:5531 | 71067
. . MDC_DEV_PDIALY_MACH_STATE_CHAN | Subsystem ¦ Machine State ¦ ¦ Dialysis, Peritoneal | Subsystem | Machine State |  | Dialysis, Peritoneal |  | Machine state | Machine state channel | All | All: All; O: O; O: X |  |  |  |  | 1 |  | MVC(3) | 1 | 5535 | 1:5535 | 71071
. . .MDC_ATTR_ALERT_CODE | ErrorEvent (numeric) ¦ Safety Systen, ‹Notifications› ¦ Machine State ¦ Dialysis, Peritoneal | ErrorEvent (numeric) | Safety Systen, ‹Notifications› | Machine State | Dialysis, Peritoneal |  | Alert ID | The number associated with the current alert. If no alert is active, then this object is not reported. | All | All | String | XXXX | N/A | -- | O | X | -- |  | -- | 0 to 9999 | 10 | 0 |  | 1(0) | 1 | 2953 | 1:2953 | 68489
. . . MDC_ATTR_ALERT_TEXT | ErrorEvent (string) ¦ Safety Systen, ‹Notifications› ¦ Machine State ¦ Dialysis, Peritoneal | ErrorEvent (string) | Safety Systen, ‹Notifications› | Machine State | Dialysis, Peritoneal |  | Alert Text | The text name of the current alert. If no alert is active, then this object is not reported. In most case this object will contain the text shown to the user. | All | All | String | N/A | N/A | -- | O | X | -- |  | -- | N/A | Venous Air Detected | 0 |  | 1(0) | 1 | 3010 | 1:3010 | 68546
. . . MDC_PDIALY_AIR_SENSOR_STATE | Air ¦ System ¦ Machine State ¦ Dialysis, Peritoneal | Air | System | Machine State | Dialysis, Peritoneal |  | Air sensor state | The state of the air sensor. | All | All | Enumeration | Alphanumeric | N/A | -- | O | X | PD_TBL_11 | AIR FLUID UNKNOWN | Air Fluid Unknown |  |  | 1 |  | 1(0) | 2 | 27613 | 2:27613 | 158685
. . . MDC_PDIALY_PUMP_SPEED | Speed ¦ Pump ¦ Machine State ¦ Dialysis, Peritoneal | Speed | Pump | Machine State | Dialysis, Peritoneal |  | Pump Speed | The speed of the pump. | All | All | Numeric | XXX.X | ml/min | ml/min | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28188 | 2:28188 | 159260
. . . MDC_PDIALY_PUMP_PRESSURE | Pressure ¦ Pump ¦ Machine State ¦ Dialysis, Peritoneal | Pressure | Pump | Machine State | Dialysis, Peritoneal |  | Pump Pressure | The pressure at the pump outlet. | All | All | Numeric | XXX.X | mmHg | mm[Hg] | O | X | -- |  | -- |  |  | 1 |  | MMM(0) | 2 | 28192 | 2:28192 | 159264
. . . MDC_PDIALY_TEMP_DIALYSATE_IN | Temperature, In ¦ Dialysate ¦ Machine State ¦ Dialysis, Peritoneal | Temperature, In | Dialysate | Machine State | Dialysis, Peritoneal | tech high low thr | Dialysate Temperature Sensor Input | Dialysate temperature at the inlet. | All | All | Numeric | XX.X | ˚C | Cel | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28196 | 2:28196 | 159268
. . . MDC_PDIALY_TEMP_DIALYSATE_OUT | Temperature, Out ¦ Dialysate ¦ Machine State ¦ Dialysis, Peritoneal | Temperature, Out | Dialysate | Machine State | Dialysis, Peritoneal | tech high low thr | Dialysate Temperature Sensor Output | Dialysate temperature at the outlet. | All | All | Numeric | XX.X | ˚C | Cel | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28200 | 2:28200 | 159272
. . . MDC_PDIALY_TEMP_HEATER_BAG | Temperature ¦ Heater Bag ¦ Machine State ¦ Dialysis, Peritoneal | Temperature | Heater Bag | Machine State | Dialysis, Peritoneal |  | Dialysate Heater Bag Temperature Sensor | Dialysate temperature in the heater bag. | All | All | Numeric | XX.X | ˚C | Cel | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28204 | 2:28204 | 159276
. . . MDC_TEMP_DEVICE_INTERNAL | Temperature ¦ Device, Internal ¦ Machine State ¦ Dialysis, Peritoneal | Temperature | Device, Internal | Machine State | Dialysis, Peritoneal |  | Device Internal Temperature Monitoring | Temperature inside the cycler. | All | All | Numeric | XX.X | ˚C | Cel | O | X | N/A |  | N/A |  |  | 0 |  | MMM(0) | 2 | 57344 | 2:57344 | 188416
. . . MDC_PDIALY_AIR_IN_VOL | Air, In ¦ Patient ¦ Machine State ¦ Dialysis, Peritoneal | Air, In | Patient | Machine State | Dialysis, Peritoneal |  | Inlet Air Volume | Measurement of the air going to the patient. | All | All | Numeric | XXX | mL | mL | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28208 | 2:28208 | 159280
. . . MDC_PDIALY_AIR_OUT_VOL | Air, Out ¦ Patient ¦ Machine State ¦ Dialysis, Peritoneal | Air, Out | Patient | Machine State | Dialysis, Peritoneal |  | Outlet Air Volume | Measurement of the air coming from the patient. | All | All | Numeric | XXX | mL | mL | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28212 | 2:28212 | 159284
. . . MDC_PDIALY_FLUID_LEAK_SENSOR_STATE | Leak ¦ ¦ Machine State ¦ Dialysis, Peritoneal | Leak |  | Machine State | Dialysis, Peritoneal |  | Leak Sensor(s) | The state of the cycler’s fluid leak sensor. | All | All | Enumeration | Alphanumeric | N/A | -- | O | X | N/A | LEAK NOLEAK UNKNOWN | N/A |  |  | 1 |  | 1(0) | 2 | 27614 | 2:27614 | 158686
. . . MDC_ATTR_AC_MAINS_RMS | Voltage ¦ AC Mains ¦ Machine State ¦ Dialysis, Peritoneal | Voltage | AC Mains | Machine State | Dialysis, Peritoneal |  | Mains Power Monitoring | Monitors AC line in. | All | All | Numeric | XXX.X | V | V | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28216 | 2:28216 | 159288
. . . MDC_PDIALY_VOLTAGE_MOTOR | Voltage ¦ Pump ¦ Machine State ¦ Dialysis, Peritoneal | Voltage | Pump | Machine State | Dialysis, Peritoneal |  | Motor Voltage Monitoring | Motor Voltage Monitoring | All | All | Numeric | XXX.X | V | V | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28220 | 2:28220 | 159292
. . . MDC_PDIALY_VOLTAGE_HEATER | Voltage ¦ Heater Bag ¦ Machine State ¦ Dialysis, Peritoneal | Voltage | Heater Bag | Machine State | Dialysis, Peritoneal |  | Heater Input Voltage | The heater input voltage level is monitored. | All | All | Numeric | XXX.X | V | V | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28224 | 2:28224 | 159296
. . . MDC_ATTR_BATT_VOLTAGE | Voltage ¦ Battery ¦ Machine State ¦ Dialysis, Peritoneal | Voltage | Battery | Machine State | Dialysis, Peritoneal |  | Battery Charge Level Monitoring | The battery voltage level is monitored. | All | All | Numeric | XX.X | V | V | O | X | N/A |  | N/A |  |  | 0 |  | 1(0) | 1 | 2485 | 1:2485 | 68021
. . . MDC_ATTR_POWER_SUPPLY_VOLTAGE | Voltage ¦ Power Supply ¦ Machine State ¦ Dialysis, Peritoneal | Voltage | Power Supply | Machine State | Dialysis, Peritoneal |  | Voltage Monitoring | The cycler monitors power supply voltages. | All | All | Numeric | XX.X | V | V | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28228 | 2:28228 | 159300
. . . MDC_PDIALY_PRESSURE_PNEUMATIC | Pressure ¦ Compressed Air ¦ Machine State ¦ Dialysis, Peritoneal | Pressure | Compressed Air | Machine State | Dialysis, Peritoneal |  | Pneumatic: Compressed air pressure | Air pressure is monitored. Compressed air is positive pressure. | All | All | Numeric | XXXX | mmHg | mm[Hg] | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28232 | 2:28232 | 159304
. . . MDC_PDIALY_PRESSURE_VACUUM | Pressure ¦ Vacuum ¦ Machine State ¦ Dialysis, Peritoneal | Pressure | Vacuum | Machine State | Dialysis, Peritoneal |  | Pneumatic: Vacuum pressure | The current pressure from the vacuum pump. Vacuum is negative pressure. | All | All | Numeric | XXXX | mmHg | mm[Hg] | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28236 | 2:28236 | 159308
. . . MDC_PDIALY_PRESSURE_HYDRAULIC | Pressure ¦ Pneumatics ¦ Machine State ¦ Dialysis, Peritoneal | Pressure | Pneumatics | Machine State | Dialysis, Peritoneal |  | Hydraulic Fluid Pressure | The current pressure from the hydraulic pump. | All | All | Numeric | XXXX | mmHg | mm[Hg] | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28240 | 2:28240 | 159312
. . . MDC_PDIALY_VOLUME_SUPPLY_BAG | Volume ¦ Supply Bag ¦ Machine State ¦ Dialysis, Peritoneal | Volume | Supply Bag | Machine State | Dialysis, Peritoneal |  | Dialysate Supply Bag Volume | The volume remaining in the current the supply bag. Note that for systems with more than one supply bag this value will have a saw tooth pattern resulting from the value going from 0 back up to the volume of a full bag when a bag switch occurs. | All | All | Numeric | X.XXX | L | L | O | X | N/A |  |  |  |  | 1 |  | MMM(0) | 2 | 28244 | 2:28244 | 159316
. . . MDC_PDIALY_VOLUME_HEATER_BAG | Volume ¦ Heater Bag ¦ Machine State ¦ Dialysis, Peritoneal | Volume | Heater Bag | Machine State | Dialysis, Peritoneal |  | Dialysate Heater Bag Volume | The current volume of the heater bag. | All | All | Numeric | X.XXX | L | L | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 28248 | 2:28248 | 159320
. . . MDC_PDIALY_WEIGHT_HEATER_BAG | Weight ¦ Heater Bag ¦ Machine State ¦ Dialysis, Peritoneal | Weight | Heater Bag | Machine State | Dialysis, Peritoneal |  | Heater Bag Weight | Weight of the heater bag | All | All | Numeric | XXXX.X | g | g | O | X | -- |  | -- |  |  | 1 |  | MMM(0) | 2 | 28252 | 2:28252 | 159324
. MDC_DEV_PDIALY_PATIENT_VMD | Device ¦ Patient ¦ ¦ Dialysis, Peritoneal | Device | Patient |  | Dialysis, Peritoneal |  | Patient "Device" |  | Intradialytic | All: All; O: O; O: O |  |  |  |  | 1 |  | MVC(2) | 1 | 5538 | 1:5538 | 71074
. . MDC_DEV_PATIENT_INFO | Subsystem ¦ Patient ¦ ¦ Dialysis, Peritoneal | Subsystem | Patient |  | Dialysis, Peritoneal |  | Patient channel | A channel to hold information about the patient. | Intradialytic | All: All; O: O; O: O |  |  |  |  | 1 |  | MVC(3) | 1 | 5543 | 1:5543 | 71079
. . . MDC_PDIALY_PATIENT_TRANSPORT_TYPE | Transport Type ¦ ¦ Patient ¦ Dialysis, Peritoneal | Transport Type |  | Patient | Dialysis, Peritoneal |  | Transport Type | The rate at which small molecules cross the peritoneal membrane. | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | O | O | PD_TBL_08 | LOW MEDLOW MEDHIGH HIGH | Low Medium Low Medium High High |  |  | 1 |  | 1(0) | 2 | 27615 | 2:27615 | 158687
. . . MDC_PDIALY_PATIENT_CATHETER_TYPE | Catheter Type ¦ ¦ Patient ¦ Dialysis, Peritoneal | Catheter Type |  | Patient | Dialysis, Peritoneal |  | Catheter Type | Catheter Type | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | O | O | PD_TBL_09 | STRAIGHT-TENCK OREO-ZELL TWH COILED-TENCK STRAIGHT-SWAN-TENCK COILED-SWAN-TENCK STRAIGHT-MO-SWAN COLIED-MO-SWAN MON-POP VINCENZA DO-PAOLA VALLI ASH-ADV CRUZ OTHER | Straight Tenckhoff Oreopoulos-Zellerman Toronto Western Hospital Coiled Tenckhoff Straight Swan-neck Tenckhoff Coiled Swan-neck Tenckhoff Straight Missouri Swan-neck Coiled Missouri Swan-neck Moncrief-Popovich Vincenza Do Paolo Valli (Balloon) Ash Advantage (T-Fluted) Cruz (Pail-Handle) Other/Unknown |  |  | 1 |  | 1(0) | 2 | 27616 | 2:27616 | 158688
. . . MDC_PDIALY_PATIENT_CATHETER_DATE | Catheter Date ¦ ¦ Patient ¦ Dialysis, Peritoneal | Catheter Date |  | Patient | Dialysis, Peritoneal |  | Catheter Placement date | Catheter Placement date | Intradialytic | All | Date | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 2 | 27617 | 2:27617 | 158689
. . . MDC_PDIALY_PAT_TARGET_WT | Weight ¦ Target ¦ Patient ¦ Dialysis, Peritoneal | Weight | Target | Patient | Dialysis, Peritoneal |  | Target Weight | The patient's target weight. This is captured for information purposes only and is not used by the dialysis machine when determining treatment settings. | Intradialytic | All | Numeric | XXX.X | Kg | kg | O | O | N/A |  | N/A |  | 75.0 | 1 |  | 1(0) | 2 | 27618 | 2:27618 | 158690
. MDC_DEV_PRESS_BLD_NONINV_VMD | Device ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Device | Pressure, noninvasive | Blood | CVS |  | Non-invasive blood pressure | Instrument for the non-invasive measurement of blood pressure. | Intradialytic | All: All; O: O; O: O |  |  |  |  | 1 |  | MVC (2) | 1 | 5150 | 1:5150 | 70686
. . MDC_DEV_PRESS_BLD_NONINV_CHAN | Subsystem ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Subsystem | Pressure, noninvasive | Blood | CVS |  |  |  | Intradialytic | All: All; O: O; O: O |  |  |  |  | 1 |  | MVC (3) | 1 | 5151 | 1:5151 | 70687
. . . MDC_DIALY_BLD_MEAS_PHASE | Measurement Phase ¦ Pressure, noninvasive ¦ Blood ¦ Dialysis, Peritoneal | Measurement Phase | Pressure, noninvasive | Blood | Dialysis, Peritoneal |  | Blood Pressure Measurement Phase | The phase of the treatment when the patient should perform a blood pressure measurement. | Intradialytic | All | String |  | N/A | -- | O | O | PD_TBL_10 | NONE PRE POST BOTH | None Pre Post Both | N/A | BOTH | 1 |  | 1(0) | 2 | 27619 | 2:27619 | 158691
. . . MDC_DIALY_PUL_MEAS_PHASE | Measurement Phase ¦ Pulse, noninvasive ¦ Blood ¦ Dialysis, Peritoneal | Measurement Phase | Pulse, noninvasive | Blood | Dialysis, Peritoneal |  | Pulse Measurement Phase | The phase of the treatment when the patient should perform a pulse measurement. | Intradialytic | All | String |  | N/A | -- | O | O | PD_TBL_10 | NONE PRE POST BOTH | None Pre Post Both | N/A | BOTH | 1 |  | 1(0) | 2 | 27620 | 2:27620 | 158692
. . . MDC_ATTR_TIME_PD_MSMT | Duration ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Duration | Pressure, noninvasive | Blood | CVS |  | Blood Pressure Measurement Interval Setting | Time allowed to elapse between automatic blood pressure measurements. A value of zero indicates that automatic measurements are not being taken. | Intradialytic | All | Numeric | XXX | minutes | min | O | O |  |  | -- | 0 to 999 | 30 | 1 |  | 1(0) | 1 | 2443 | 1:2443 | 67979
. . . MDC_ATTR_PT_BODY_POSN | Body Position (enumeration) ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Body Position (enumeration) | Pressure, noninvasive | Blood | CVS |  | Patient body position | Patient body position at the time of the {blood pressure} measurement. | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | O | O | _TBL_04 – Patient Position | SITTING STANDING SUPINE | SITTING STANDING SUPINE | N/A | SUPINE | 1 |  | 1(0) | 1 | 2599 | 1:2599 | 68135
. . . MDC_PRESS_BLD_NONINV_DIA | Pressure ¦ Noninvasive, Diastolic ¦ Blood ¦ CVS | Pressure | Noninvasive, Diastolic | Blood | CVS | phys high low thr | Diastolic Pressure | Minimum arterial pressure during relaxation and dilatation of the ventricles of the heart when the ventricles fill with blood | Intradialytic | All | Numeric | XXX | mmHg | mm[Hg] | O | X |  |  | -- | 0 to 999 | 80 | 1 |  | SDM (2) | 2 | 18950 | 2:18950 | 150022
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Noninvasive, Diastolic ¦ Blood ¦ Dialysis, Peritoneal | Entry Range | Noninvasive, Diastolic | Blood | Dialysis, Peritoneal |  | Diastolic Pressure Entry Limits | The minimum and maximum value that the user can enter for a diastolic pressure. | Intradialytic | All | Numeric Range | XXX^XXX | mmHg | mm[Hg] | O | O |  |  |  | 0 to 999 | 0^999 | 1 |  | 1(0) | 1 | 2464 | 1:2464 | 68000
. . . MDC_PULS_RATE_NON_INV | Rate ¦ Noninvasive, Pulse ¦ Blood ¦ CVS | Rate | Noninvasive, Pulse | Blood | CVS | phys high low thr | Heart Rate | Number of heart beats per minute | Intradialytic | All | Numeric | XXX | BPM | {beats}/min | O | X |  |  | -- | 0 to 999 | 70 | 0 |  | RCE (2),, | 2 | 18474 | 2:18474 | 149546
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Noninvasive, Pulse ¦ Blood ¦ Dialysis, Peritoneal | Entry Range | Noninvasive, Pulse | Blood | Dialysis, Peritoneal |  | Pulse Entry Limits | The minimum and maximum value that the user can enter for a pulse measurement. | Intradialytic | All | Numeric Range | XXX^XXX | BPM | {beats}/min | O | O |  |  |  | 0 to 999 | 20^200 | 0 |  | 1(0) | 1 | 2464 | 1:2464 | 68000
. . . MDC_PRESS_BLD_NONINV_MEAN | Pressure ¦ Noninvasive, Mean ¦ Blood ¦ CVS | Pressure | Noninvasive, Mean | Blood | CVS |  | Mean Arterial Pressure | Average pressure in a patient's arteries during one cardiac cycle | Intradialytic | All | Numeric | XXX | mmHg | mm[Hg] | O | X |  |  | -- | 0 to 999 | 100 | 0 |  | SDM (3) | 2 | 18951 | 2:18951 | 150023
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Noninvasive, Mean ¦ Blood ¦ Dialysis, Peritoneal | Entry Range | Noninvasive, Mean | Blood | Dialysis, Peritoneal |  | Mean Arterial Pressure Entry Limits | The minimum and maximum value that the user can enter for a mean arterial pressure. | Intradialytic | All | Numeric Range | XXX^XXX | mmHg | mm[Hg] | O | O |  |  |  | 0 to 999 | 0^999 | 0 |  | 1(0) | 1 | 2464 | 1:2464 | 68000
. . . MDC_PRESS_BLD_NONINV_SYS | Pressure ¦ Noninvasive, Systolic ¦ Blood ¦ CVS | Pressure | Noninvasive, Systolic | Blood | CVS | phys high low thr | Systolic Pressure | Maximum arterial pressure during contraction of the left ventricle of the heart | Intradialytic | All | Numeric | XXX | mmHg | mm[Hg] | O | X |  |  | -- | 0 to 999 | 120 | 0 |  | SDM (1) | 2 | 18949 | 2:18949 | 150021
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Noninvasive, Systolic ¦ Blood ¦ Dialysis, Peritoneal | Entry Range | Noninvasive, Systolic | Blood | Dialysis, Peritoneal |  | Systolic Pressure Entry Limits | The minimum and maximum value that the user can enter for a systolic pressure. | Intradialytic | All | Numeric Range | XXX^XXX | mmHg | mm[Hg] | O | O |  |  |  | 0 to 999 | 0^999 | 0 |  | 1(0) | 1 | 2464 | 1:2464 | 68000
. MDC_DEV_SPEC_PROFILE_GLUCOSE | Device ¦ Glucose ¦ Blood ¦ CVS | Device | Glucose | Blood | CVS |  | Glucose "Device" |  | Intradialytic | All: All; O: O; O: O |  |  |  |  | 0 |  | 1(0) | 8 | 4113 | 8:4113 | 528401
. . MDC_DEV_CHAN | Subsystem ¦ Glucose ¦ Blood ¦ CVS | Subsystem | Glucose | Blood | CVS |  | Glucose Channel | A channel to hold information about the patient's glucose. | Intradialytic | All: All; O: O; O: O |  |  |  |  | 0 |  | MVC(3) | 1 | 4099 | 1:4099 | 69635
. . . MDC_DIALY_PAT_GLU_MEAS_PHASE | Measurement Phase ¦ Glucose ¦ Blood ¦ Dialysis, Peritoneal | Measurement Phase | Glucose | Blood | Dialysis, Peritoneal |  | Glucose Measurement Phase | The phase of the treatment when the patient should perform a glocuse measurement. | Intradialytic | All | String |  | N/A | -- | O | O | PD_TBL_10 | NONE PRE POST BOTH | None Pre Post Both |  |  | 1 |  | 1(0) | 2 | 27621 | 2:27621 | 158693
. . . MDC_CONC_GLU_CAPILLARY_PLASMA | Measurement ¦ Glucose ¦ Blood ¦ Dialysis, Peritoneal | Measurement | Glucose | Blood | Dialysis, Peritoneal |  | Blood Glucose | Blood glucose at the start of treatment | Intradialytic | All | Numeric | XX.X | mmol/L | mmol/L | O | X | N/A |  | N/A |  |  | 1 |  | MMM(0) | 2 | 29116 | 2:29116 | 160188
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Glucose ¦ Blood ¦ Dialysis, Peritoneal | Entry Range | Glucose | Blood | Dialysis, Peritoneal |  | Glucose Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XX.X^XX.X | mmol/L | mmol/L | O | O | N/A |  | N/A |  |  | 1 |  | 1(0) | 1 | 2464 | 1:2464 | 68000
. MDC_DEV_SPEC_PROFILE_SCALE | Device ¦ Weight ¦ Scale ¦ CVS | Device | Weight | Scale | CVS |  | Scale "Device" |  | Intradialytic | All: All; O: O; O: O |  |  |  |  | 1 |  | 1(0) | 8 | 4111 | 8:4111 | 528399
. . MDC_DEV_CHAN | Subsystem ¦ Weight ¦ Scale ¦ CVS | Subsystem | Weight | Scale | CVS |  | Weight Channel | A channel to hold information about the patient's weight. | Intradialytic | All: All; O: O; O: O |  |  |  |  | 1 |  | MVC(3) | 1 | 4099 | 1:4099 | 69635
. . . MDC_DIALY_PAT_WT_MEAS_PHASE | Measurement Phase ¦ Weight ¦ Patient ¦ Dialysis, Peritoneal | Measurement Phase | Weight | Patient | Dialysis, Peritoneal |  | Weight Measurement Phase | The phase of the treatment when the patient should perform a weight measurement. | Intradialytic | All | String |  | N/A | -- | O | O | PD_TBL_10 | NONE PRE POST BOTH | None Pre Post Both | N/A | PRE | 1 |  | 1(0) | 2 | 27622 | 2:27622 | 158694
. . . MDC_MASS_BODY_ACTUAL | Measurement ¦ Weight ¦ Patient ¦ Dialysis, Peritoneal | Measurement | Weight | Patient | Dialysis, Peritoneal |  | Patient Weight | Weight of patient at beginning or end of treatment | Intradialytic | All | Numeric | XXX.X | Kg | kg | O | X | N/A |  | N/A | 0 to 999.9 | 75.9 | 1 |  | MMM | 2 | 57664 | 2:57664 | 188736
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Weight ¦ Patient ¦ Dialysis, Peritoneal | Entry Range | Weight | Patient | Dialysis, Peritoneal |  | Weight Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XXX.X^XXX.X | Kg | kg | O | O | N/A |  | N/A | 0 to 999.9 | 60.0 | 1 |  | 1(0) | 1 | 2464 | 1:2464 | 68000
. MDC_DEV_SPEC_PROFILE_TEMP | Device ¦ Temperature ¦ Thermometer ¦ CVS | Device | Temperature | Thermometer | CVS |  | Thermometer "Device" |  | Intradialytic | All: All; O: O; O: O |  |  |  |  | 1 |  | 1(0) | 8 | 4104 | 8:4104 | 528392
. . MDC_DEV_CHAN | Subsystem ¦ Temperature ¦ Thermometer ¦ CVS | Subsystem | Temperature | Thermometer | CVS |  | Temperature Channel | A channel to hold information about the patient's temperature. | Intradialytic | All: All; O: O; O: O |  |  |  |  | 1 |  | MVC(3) | 1 | 4099 | 1:4099 | 69635
. . . MDC_DIALY_PAT_TEMP_MEAS_PHASE | Measurement Phase ¦ Temperature ¦ Patient ¦ Dialysis, Peritoneal | Measurement Phase | Temperature | Patient | Dialysis, Peritoneal |  | Temperature Measurement Phase | The phase of the treatment when the patient should perform a temperature measurement. | Intradialytic | All | String |  | N/A | -- | O | O | PD_TBL_10 | NONE PRE POST BOTH | None Pre Post Both | N/A | NONE | 1 |  | 1(0) | 2 | 27623 | 2:27623 | 158695
. . . MDC_TEMP_TYMP | Measurement ¦ Temperature, Tympanic ¦ Patient ¦ Dialysis | Measurement | Temperature, Tympanic | Patient | Dialysis | phys high thr | Patient Temperature | Patient's temperature when measured at the ear drum. | Intradialytic | All | Numeric | XX.X | oC | Cel | O | X | N/A |  | N/A | 0 to 99.9 | 39.0 | 0 |  | MMM(0) | 2 | 19320 | 2:19320 | 150392
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Temperature, Tympanic ¦ Patient ¦ Dialysis, Peritoneal | Entry Range | Temperature, Tympanic | Patient | Dialysis, Peritoneal |  | Temperature Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XX.X^XX.X | oC | Cel | O | O | N/A |  | N/A | 0 to 99.9 | 35.0^41.1 | 0 |  | 1(0) | 1 | 2464 | 1:2464 | 68000
. . . MDC_TEMP_ORAL | Measurement ¦ Temperature, Oral ¦ Patient ¦ Dialysis | Measurement | Temperature, Oral | Patient | Dialysis | phys high thr | Patient Temperature | Patient's temperature when measured orally. | Intradialytic | All | Numeric | XX.X | oC | Cel | O | X | N/A |  | N/A | 0 to 99.9 | 39.0 | 0 |  | MMM(0) | 2 | 57352 | 2:57352 | 188424
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Temperature, Oral ¦ Patient ¦ Dialysis, Peritoneal | Entry Range | Temperature, Oral | Patient | Dialysis, Peritoneal |  | Temperature Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XX.X^XX.X | oC | Cel | O | O | N/A |  | N/A | 0 to 99.9 | 35.0^41.1 | 0 |  | 1(0) | 1 | 2464 | 1:2464 | 68000
. . . MDC_TEMP_BODY | Measurement ¦ Temperature, Body ¦ Patient ¦ Dialysis | Measurement | Temperature, Body | Patient | Dialysis | phys high thr | Patient Temperature | Patient's temperature when measured at an unknown location | Intradialytic | All | Numeric | XX.X | oC | Cel | O | X | N/A |  | N/A | 0 to 99.9 | 39.0 | 0 |  | MMM(0) | 2 | 19292 | 2:19292 | 150364
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Temperature, Body ¦ Patient ¦ Dialysis, Peritoneal | Entry Range | Temperature, Body | Patient | Dialysis, Peritoneal |  | Temperature Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XX.X^XX.X | oC | Cel | O | O | N/A |  | N/A | 0 to 99.9 | 35.0^41.1 | 0 |  | 1(0) | 1 | 2464 | 1:2464 | 68000
. . . MDC_TEMP_AXILLA | Measurement ¦ Temperature, Axillary ¦ Patient ¦ Dialysis | Measurement | Temperature, Axillary | Patient | Dialysis | phys high thr | Patient Temperature | Patient's temperature when measured at the arm pit. | Intradialytic | All | Numeric | XX.X | oC | Cel | O | X | N/A |  | N/A | 0 to 99.9 | 39.0 | 0 |  | MMM(0) | 2 | 57380 | 2:57380 | 188452
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Temperature, Axillary ¦ Patient ¦ Dialysis, Peritoneal | Entry Range | Temperature, Axillary | Patient | Dialysis, Peritoneal |  | Temperature Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XX.X^XX.X | oC | Cel | O | O | N/A |  | N/A | 0 to 99.9 | 35.0^41.1 | 0 |  | 1(0) | 1 | 2464 | 1:2464 | 68000

## Events
Logical Source | Source | Event | SysName | SysNameBase | SysName1stDesc | SysName2ndDesc | SysName3rdDesc | Alert Type | Dialysis Common Name | Definition | Usage | gp | s | Disc | PART | CODE10 | PartCode | CF_CODE10
‹vendor-text› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_INTERNAL_SYSTEM_FAULT | ErrorEvent ¦ System, Internal, ‹Fault› ¦ FunctionalDisturbance ¦ Device | ErrorEvent | System, Internal, ‹Fault› | FunctionalDisturbance | Device | tech |  | Manufacturer specific system fault. | O | 0 |  | EVT(0) | 3 | 1178 | 3::1178 | 197786
‹vendor-text› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_MSMT_OUT_OF_RANGE | ErrorEvent ¦ Measurement, ‹Out of range› ¦ Processing ¦ Device | ErrorEvent | Measurement, ‹Out of range› | Processing | Device | tech |  | Manufacturer specific measurement out of range. | O | 0 |  | EVT(0) | 3 | 1242 | 3::1242 | 197850
‹vendor-text› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_SELFTEST_FAILURE | ErrorEvent ¦ Selftest, ‹Failure› ¦ Functional Disturbance ¦ Device | ErrorEvent | Selftest, ‹Failure› | Functional Disturbance | Device | tech |  | Manufacturer specific selftest failure. | O | 0 |  | EVT(0) | 3 | 1608 | 3::1608 | 198216
‹vendor-text› | MDC_DEV_PDIALY_VMD | MDC_EVT_ADVIS_USER_MESSAGE | Message, User ¦ System ¦ System ¦ Dialysis, Peritoneal | Message, User | System | System | Dialysis, Peritoneal | tech |  | Machine specific advisory | O | 1 | 4 | EVT(0) | 3 | 6900 | 3::6900 | 203508
‹vendor-text› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_PATIENT_ID_INVALID | Invalid ¦ Patient ID ¦ System ¦ Patient | Invalid | Patient ID | System | Patient | tech |  | Invalid Patient ID | O | 1 | 0 | EVT(0) | 3 | 1754 | 3::1754 | 198362
‹vendor-text› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_TREATMENT_ERR | General Error ¦ Treatment ¦ System ¦ Patient | General Error | Treatment | System | Patient | tech |  | An error occurred during treatment | O | 1 | 0 | EVT(0) | 3 | 1756 | 3::1756 | 198364
‹vendor-text› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_UNRECOV_TREATMENT_ERR | Unrecoverable ¦ Treatment ¦ System ¦ Patient | Unrecoverable | Treatment | System | Patient | tech |  | An unrecoverable error occurred during treatment | O | 1 | 0 | EVT(0) | 3 | 1758 | 3::1758 | 198366
MDC_ATTR_PT_WEIGHT | MDC_MASS_BODY_ACTUAL | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Setting, ‹CheckingNecessary› ¦ FunctionalDisturbance ¦ Device | Advisory | Setting, ‹CheckingNecessary› | FunctionalDisturbance | Device | tech |  | The user should check the patient weight. | O | 0 |  | EVT(0) | 3 | 6838 | 3::6838 | 203446
‹drain-line› | MDC_DEV_PDIALY_DRAIN_LINE_CHAN | MDC_EVT_FLUID_FLOW_BLOCKED | Blockage ¦ Fluid Flow ¦ System ¦ Device, Fluid Line or Reservoir | Blockage | Fluid Flow | System | Device, Fluid Line or Reservoir | tech |  | A line is blocked; checking is required. | O | 1 | 0 | EVT(0) | 3 | 1762 | 3::1762 | 198370
‹drain-line› | MDC_DEV_PDIALY_DRAIN_LINE_CHAN | MDC_EVT_FLUID_FLOW_BLOCKED_CHK_REQD | Blockage Check ¦ Fluid ¦ System ¦ Device, Fluid Line or Reservoir | Blockage Check | Fluid | System | Device, Fluid Line or Reservoir | tech |  | Slow flow in drain line should be checked by user. | O | 1 | 0 | EVT(0) | 3 | 1764 | 3::1764 | 198372
‹patient-and-or-drain-line› | MDC_DEV_PDIALY_DRAIN_LINE_CHAN | MDC_EVT_FLUID_FLOW_BLOCKED_OUTFLOW | Blockage ¦ Fluid Outflow ¦ System ¦ Device, Fluid Line | Blockage | Fluid Outflow | System | Device, Fluid Line | tech |  | Bloced flow in drain line. | O | 1 | 0 | EVT(0) | 3 | 1766 | 3::1766 | 198374
‹drain-line› | MDC_DEV_PDIALY_DRAIN_LINE_CHAN | MDC_EVT_FLUID_FLOW_SLOW | Slow Flow ¦ Fluid Flow ¦ System ¦ Device, Fluid Line or Reservoir | Slow Flow | Fluid Flow | System | Device, Fluid Line or Reservoir | tech |  | Slow flow in drain line. | O | 1 | 0 | EVT(0) | 3 | 1768 | 3::1768 | 198376
‹drain-line› | MDC_DEV_PDIALY_DRAIN_LINE_CHAN | MDC_EVT_FLUID_LINE_CHK | Check ¦ Line ¦ Fluid Line ¦ Device, Fluid Line or Reservoir | Check | Line | Fluid Line | Device, Fluid Line or Reservoir | tech |  | Check drain line. | O | 1 | 0 | EVT(0) | 3 | 1770 | 3::1770 | 198378
‹heater-line› | MDC_DEV_PDIALY_HEATER_LINE_CHAN | MDC_EVT_FLUID_FLOW_BLOCKED | Blockage ¦ Fluid Flow ¦ System ¦ Device, Fluid Line or Reservoir | Blockage | Fluid Flow | System | Device, Fluid Line or Reservoir | tech |  | A line is blocked | O | 1 | 0 | EVT(0) | 3 | 1762 | 3::1762 | 198370
‹heater-line› | MDC_DEV_PDIALY_HEATER_LINE_CHAN | MDC_EVT_FLUID_FLOW_SLOW | Slow Flow ¦ Fluid Flow ¦ System ¦ Device, Fluid Line or Reservoir | Slow Flow | Fluid Flow | System | Device, Fluid Line or Reservoir | tech |  | Slow flow in heater line. | O | 1 | 0 | EVT(0) | 3 | 1768 | 3::1768 | 198376
‹heater-bag-or-line› | MDC_DEV_PDIALY_HEATER_LINE_CHAN | MDC_EVT_PDIALY_HEATER_BAG_POSN_ERR_OR_LINE_CLAMPED_CHK_REQD | Check ¦ Heater Bag ¦ System ¦ Dialysis, Peritoneal | Check | Heater Bag | System | Dialysis, Peritoneal | tech |  | Verify that the heater bag is on the heater tray and the heater line is not clamped | O | 1 | 0 | EVT(0) | 3 | 1772 | 3::1772 | 198380
‹heater-line› | MDC_DEV_PDIALY_HEATER_LINE_CHAN | MDC_EVT_FLUID_LINE_CHK | Check ¦ Line ¦ Fluid Line ¦ Device, Fluid Line or Reservoir | Check | Line | Fluid Line | Device, Fluid Line or Reservoir | tech |  | Check heater line. | O | 1 | 0 | EVT(0) | 3 | 1770 | 3::1770 | 198378
‹program-execution› | MDC_DEV_PDIALY_MACH_STATE_CHAN | MDC_EVT_PROGRAM_SYNC_ERROR | Error ¦ Program execution, Synchonization ¦ System ¦ Device | Error | Program execution, Synchonization | System | Device | tech |  | Error during monitoring of program execution synchronisation | O | 1 | 0 | EVT(0) | 3 | 1774 | 3::1774 | 198382
‹mds› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_AC_POWER_FAIL | StatusEvent ¦ AC Power, ‹Failure› ¦ FunctionalDisturbance ¦ PowerSupply | StatusEvent | AC Power, ‹Failure› | FunctionalDisturbance | PowerSupply | tech |  | AC Power fail (or short circuit in one of the devices connected to an auxiliary outlet) | O | 0 |  | EVT(0) | 3 | 838 | 3::838 | 197446
‹mds› ‹mds-or-vmd› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_BATT_LO | ErrorEvent ¦ Battery, ‹Low› ¦ FunctionalDisturbance ¦ Device | ErrorEvent | Battery, ‹Low› | FunctionalDisturbance | Device | tech | Low Battery | {device or sensor} Battery low | O | 0 |  | EVT(0) | 3 | 194 | 3::194 | 196802
‹mds-or-vmd› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_HID_MALF | ErrorEvent ¦ HumanInterfaceDevice, ‹Malfunction› ¦ FunctionalDisturbance ¦ Device, HID | ErrorEvent | HumanInterfaceDevice, ‹Malfunction› | FunctionalDisturbance | Device, HID | tech |  | HID failure | O | 0 |  | EVT(0) | 3 | 1624 | 3::1624 | 198232
‹mds› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_INTERNAL_SYSTEM_FAULT | ErrorEvent ¦ System, Internal, ‹Fault› ¦ FunctionalDisturbance ¦ Device | ErrorEvent | System, Internal, ‹Fault› | FunctionalDisturbance | Device | tech |  | Unspecified system fault. | O | 0 |  | EVT(0) | 3 | 1178 | 3::1178 | 197786
‹mds-or-vmd› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_MSMT_OUT_OF_RANGE | ErrorEvent ¦ Measurement, ‹Out of range› ¦ Processing ¦ Device | ErrorEvent | Measurement, ‹Out of range› | Processing | Device | tech |  | Unspecified measurement out of range | O | 0 |  | EVT(0) | 3 | 1242 | 3::1242 | 197850
‹mds-or-vmd› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_SELFTEST_FAILURE | ErrorEvent ¦ Selftest, ‹Failure› ¦ Functional Disturbance ¦ Device | ErrorEvent | Selftest, ‹Failure› | Functional Disturbance | Device | tech |  | Unspecified selftest failure | O | 0 |  | EVT(0) | 3 | 1608 | 3::1608 | 198216
‹mds-or-vmd› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_UNRECOV_ERR | ErrorEvent ¦ StateMachine, ‹UnrecoverableError› ¦ Communication ¦ Device | ErrorEvent | StateMachine, ‹UnrecoverableError› | Communication | Device | tech |  | Unrecoverable state machine error in communication. | O | 0 |  | EVT(0) | 3 | 134 | 3::134 | 196742
‹mds› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_AC_POWER_RESTORED | Power Restored ¦ AC Mains ¦ System ¦ Device | Power Restored | AC Mains | System | Device | tech |  | AC power restored | O | 1 | 0 | EVT(0) | 3 | 1776 | 3::1776 | 198384
‹mds› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_DEVICE_TILTED | Tilted, off vertical axis ¦ System ¦ System ¦ Device | Tilted, off vertical axis | System | System | Device | tech |  | Device tilted | O | 1 | 0 | EVT(0) | 3 | 1778 | 3::1778 | 198386
‹mds› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_EXTERNAL_MEMORY_ERROR | Error ¦ External Memory ¦ System ¦ Device | Error | External Memory | System | Device | tech |  | External memory device error (e.g. external memory not available, cannot communicate, cannot be read, cannot be written) | O | 1 | 0 | EVT(0) | 3 | 1780 | 3::1780 | 198388
‹mds› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_SPEAKER_FAIL | ErrorEvent ¦ Speaker, ‹Failure› ¦ FunctionalDisturbance ¦ Device | ErrorEvent | Speaker, ‹Failure› | FunctionalDisturbance | Device | tech |  | The speaker power and signals are monitored to ensure the speaker is active. | O | 0 |  | EVT(0) | 3 | 1454 | 3::1454 | 198062
<vmd> | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_COMM_FAILURE | Failure ¦ Communications ¦ System ¦ Dialysis, Peritoneal | Failure | Communications | System | Dialysis, Peritoneal | tech |  | The cycler monitors the information flow to the cloud to determine that the information has been sent. | O | 1 | 0 | EVT(0) | 3 | 1782 | 3::1782 | 198390
<vmd> | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_PUMP_FAILURE | Failure ¦ Pump ¦ System ¦ Dialysis, Peritoneal | Failure | Pump | System | Dialysis, Peritoneal | tech |  | The cycler has determined that the pump is no longer functional. | O | 1 | 0 | EVT(0) | 3 | 1784 | 3::1784 | 198392
<vmd> | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_SET_POSITION_WRONG | Deviation ¦ Set Position ¦ System ¦ Dialysis, Peritoneal | Deviation | Set Position | System | Dialysis, Peritoneal | tech | Disposable Set Position Sensor | The disposable set status is monitored. | O | 1 | 0 | EVT(0) | 3 | 1786 | 3::1786 | 198394
<vmd> | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_STRAIN_GAUGE_FAILURE | Failure ¦ Strain Gauge ¦ System ¦ Dialysis, Peritoneal | Failure | Strain Gauge | System | Dialysis, Peritoneal | tech | Strain Gauge Sensor Status | The strain gauge state is monitored. | O | 1 | 0 | EVT(0) | 3 | 1788 | 3::1788 | 198396
‹mds› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_TOUCHSCREEN_FAIL | ErrorEvent ¦ Touchscreen, ‹Failure› ¦ FunctionalDisturbance ¦ Device | ErrorEvent | Touchscreen, ‹Failure› | FunctionalDisturbance | Device | tech |  | One or more elements touchscreen have failed. | O | 0 |  | EVT(0) | 3 | 1500 | 3::1500 | 198108
‹mds› | MDC_DEV_PDIALY_MACHINE_MDS | MDC_EVT_DISPLAY_FAIL | Failure ¦ Display ¦ System ¦ Device, Display | Failure | Display | System | Device, Display | tech |  | One or more elements of the display have failed. | O | 1 | 0 | EVT(0) | 3 | 1790 | 3::1790 | 198398
‹patient-line› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_PDIALY_DRAIN_COMPLICATION_CHK_REQD | Complication ¦ Drain ¦ System ¦ Dialysis, Peritoneal | Complication | Drain | System | Dialysis, Peritoneal | tech |  | Drain complication encountered: check patient and drain lines | O | 1 | 0 | EVT(0) | 3 | 1792 | 3::1792 | 198400
‹patient-line+› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_PDIALY_FILL_COMPLICATION_CHK_REQD | Complication ¦ Fill ¦ System ¦ Dialysis, Peritoneal | Complication | Fill | System | Dialysis, Peritoneal | tech |  | Fill complication encountered: check heater, patient and drain lines | O | 1 | 0 | EVT(0) | 3 | 1794 | 3::1794 | 198402
‹patient-line› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_FLUID_FLOW_BLOCKED | Blockage ¦ Fluid Flow ¦ System ¦ Device, Fluid Line or Reservoir | Blockage | Fluid Flow | System | Device, Fluid Line or Reservoir | tech |  | Tubing system blocked during inflow to patient. | O | 1 | 0 | EVT(0) | 3 | 1762 | 3::1762 | 198370
‹patient-line› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_FLUID_FLOW_BLOCKED_CHK_REQD | Blockage Check ¦ Fluid ¦ System ¦ Device, Fluid Line or Reservoir | Blockage Check | Fluid | System | Device, Fluid Line or Reservoir | tech |  | Slow flow in patient line should be checked by user. | O | 1 | 0 | EVT(0) | 3 | 1764 | 3::1764 | 198372
‹patient+› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_FLUID_FLOW_BLOCKED_INFLOW | Blockage ¦ Fluid Inflow ¦ System ¦ Device, Fluid Line | Blockage | Fluid Inflow | System | Device, Fluid Line | tech |  | Blocked flow in patient inlet line. | O | 1 | 0 | EVT(0) | 3 | 1796 | 3::1796 | 198404
‹patient-and-or-drain-line› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_FLUID_FLOW_BLOCKED_OUTFLOW | Blockage ¦ Fluid Outflow ¦ System ¦ Device, Fluid Line | Blockage | Fluid Outflow | System | Device, Fluid Line | tech |  | Blocked flow in patient outlet line. | O | 1 | 0 | EVT(0) | 3 | 1766 | 3::1766 | 198374
‹patient-line› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_FLUID_FLOW_SLOW | Slow Flow ¦ Fluid Flow ¦ System ¦ Device, Fluid Line or Reservoir | Slow Flow | Fluid Flow | System | Device, Fluid Line or Reservoir | tech |  | Slow flow in patient line. | O | 1 | 0 | EVT(0) | 3 | 1768 | 3::1768 | 198376
‹patient-line› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_FLUID_FLOW_SLOW_INFLOW | Slow Flow ¦ Fluid Inflow ¦ System ¦ Device, Fluid Line | Slow Flow | Fluid Inflow | System | Device, Fluid Line | tech |  | Slow flow in patient inlet line. | O | 1 | 0 | EVT(0) | 3 | 1798 | 3::1798 | 198406
‹patient-line› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_FLUID_FLOW_SLOW_OUTFLOW | Slow Flow ¦ Fluid Outflow ¦ System ¦ Device, Fluid Line | Slow Flow | Fluid Outflow | System | Device, Fluid Line | tech |  | Slow flow in patient outlet line. | O | 1 | 0 | EVT(0) | 3 | 1800 | 3::1800 | 198408
‹patient-line› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_FLUID_LINE_ABSENT | Missing ¦ Fluid Line ¦ System ¦ Device | Missing | Fluid Line | System | Device | tech |  | Missing or disconnected patient line. | O | 1 | 0 | EVT(0) | 3 | 1802 | 3::1802 | 198410
‹patient-line› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_FLUID_LINE_CHK | Check ¦ Line ¦ Fluid Line ¦ Device, Fluid Line or Reservoir | Check | Line | Fluid Line | Device, Fluid Line or Reservoir | tech |  | Check patient line. | O | 1 | 0 | EVT(0) | 3 | 1770 | 3::1770 | 198378
‹patient-line› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_STAT_PDIALY_PATIENT_DRAIN_PHASE | Status ¦ Drain ¦ System ¦ Dialysis, Peritoneal | Status | Drain | System | Dialysis, Peritoneal | tech |  | Patient drain phase | O | 1 | 2 | EVT(0) | 3 | 6430 | 3::6430 | 203038
‹patient-line› | MDC_DEV_PDIALY_PATIENT_LINE_CHAN | MDC_EVT_STAT_PDIALY_PATIENT_FILL_PHASE | Status ¦ Fill ¦ System ¦ Dialysis, Peritoneal | Status | Fill | System | Dialysis, Peritoneal | tech |  | Patient fill or refill phase | O | 1 | 2 | EVT(0) | 3 | 6432 | 3::6432 | 203040
<supply-line> | MDC_DEV_PDIALY_FLUID_1_CHAN MDC_DEV_PDIALY_FLUID_2_CHAN MDC_DEV_PDIALY_FLUID_3_CHAN MDC_DEV_PDIALY_FLUID_4_CHAN MDC_DEV_PDIALY_FLUID_5_CHAN MDC_DEV_PDIALY_FLUID_6_CHAN MDC_DEV_PDIALY_SUPPLY_LINE_CHAN | MDC_EVT_FLUID_FLOW_BLOCKED | Blockage ¦ Fluid Flow ¦ System ¦ Device, Fluid Line or Reservoir | Blockage | Fluid Flow | System | Device, Fluid Line or Reservoir | tech |  | Tubing system blocked during outflow from patient. | O | 1 | 0 | EVT(0) | 3 | 1762 | 3::1762 | 198370
<supply-line> | MDC_DEV_PDIALY_FLUID_1_CHAN MDC_DEV_PDIALY_FLUID_2_CHAN MDC_DEV_PDIALY_FLUID_3_CHAN MDC_DEV_PDIALY_FLUID_4_CHAN MDC_DEV_PDIALY_FLUID_5_CHAN MDC_DEV_PDIALY_FLUID_6_CHAN MDC_DEV_PDIALY_SUPPLY_LINE_CHAN | MDC_EVT_FLUID_FLOW_BLOCKED_CHK_REQD | Blockage Check ¦ Fluid ¦ System ¦ Device, Fluid Line or Reservoir | Blockage Check | Fluid | System | Device, Fluid Line or Reservoir | tech |  | Slow flow in supply line should be checked by user. | O | 1 | 0 | EVT(0) | 3 | 1764 | 3::1764 | 198372
‹supply-line› | MDC_DEV_PDIALY_FLUID_1_CHAN MDC_DEV_PDIALY_FLUID_2_CHAN MDC_DEV_PDIALY_FLUID_3_CHAN MDC_DEV_PDIALY_FLUID_4_CHAN MDC_DEV_PDIALY_FLUID_5_CHAN MDC_DEV_PDIALY_FLUID_6_CHAN MDC_DEV_PDIALY_SUPPLY_LINE_CHAN | MDC_EVT_FLUID_FLOW_SLOW | Slow Flow ¦ Fluid Flow ¦ System ¦ Device, Fluid Line or Reservoir | Slow Flow | Fluid Flow | System | Device, Fluid Line or Reservoir | tech |  | Slow flow in supply line. | O | 1 | 0 | EVT(0) | 3 | 1768 | 3::1768 | 198376
‹supply-or-final-line› | MDC_DEV_PDIALY_FLUID_1_CHAN MDC_DEV_PDIALY_FLUID_2_CHAN MDC_DEV_PDIALY_FLUID_3_CHAN MDC_DEV_PDIALY_FLUID_4_CHAN MDC_DEV_PDIALY_FLUID_5_CHAN MDC_DEV_PDIALY_FLUID_6_CHAN MDC_DEV_PDIALY_SUPPLY_LINE_CHAN | MDC_EVT_FLUID_LINE_ABSENT | Missing ¦ Fluid Line ¦ System ¦ Device | Missing | Fluid Line | System | Device | tech |  | Missing or disconnected dialysate line. | O | 1 | 0 | EVT(0) | 3 | 1802 | 3::1802 | 198410
‹supply-line› | MDC_DEV_PDIALY_FLUID_1_CHAN MDC_DEV_PDIALY_FLUID_2_CHAN MDC_DEV_PDIALY_FLUID_3_CHAN MDC_DEV_PDIALY_FLUID_4_CHAN MDC_DEV_PDIALY_FLUID_5_CHAN MDC_DEV_PDIALY_FLUID_6_CHAN MDC_DEV_PDIALY_SUPPLY_LINE_CHAN | MDC_EVT_FLUID_LINE_CHK | Check ¦ Line ¦ Fluid Line ¦ Device, Fluid Line or Reservoir | Check | Line | Fluid Line | Device, Fluid Line or Reservoir | tech |  | Check dialysate/supply line. | O | 1 | 0 | EVT(0) | 3 | 1770 | 3::1770 | 198378
‹supply-line› | MDC_DEV_PDIALY_FLUID_1_CHAN MDC_DEV_PDIALY_FLUID_2_CHAN MDC_DEV_PDIALY_FLUID_3_CHAN MDC_DEV_PDIALY_FLUID_4_CHAN MDC_DEV_PDIALY_FLUID_5_CHAN MDC_DEV_PDIALY_FLUID_6_CHAN MDC_DEV_PDIALY_SUPPLY_LINE_CHAN | MDC_EVT_PDIALY_SOLUTION_NOT_REQD | Fluid ¦ Unused Line ¦ System ¦ Dialysis, Peritoneal | Fluid | Unused Line | System | Dialysis, Peritoneal | tech |  | Dialysate solution detected on a line where it is not required by the prescription. | O | 1 | 0 | EVT(0) | 3 | 1804 | 3::1804 | 198412
num_pd_therapy_time_setting | MDC_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Setting, ‹CheckingNecessary› ¦ FunctionalDisturbance ¦ Device | Advisory | Setting, ‹CheckingNecessary› | FunctionalDisturbance | Device | tech |  | The user should check the therapy time setting. | O | 0 |  | EVT(0) | 3 | 6838 | 3::6838 | 203446
num_pd_volume_UF_total_setting | MDC_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Setting, ‹CheckingNecessary› ¦ FunctionalDisturbance ¦ Device | Advisory | Setting, ‹CheckingNecessary› | FunctionalDisturbance | Device | tech |  | The user should check the total UF setting. | O | 0 |  | EVT(0) | 3 | 6838 | 3::6838 | 203446
num_pd_volume_total_setting | MDC_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Setting, ‹CheckingNecessary› ¦ FunctionalDisturbance ¦ Device | Advisory | Setting, ‹CheckingNecessary› | FunctionalDisturbance | Device | tech |  | The user should check the total volume setting. | O | 0 |  | EVT(0) | 3 | 6838 | 3::6838 | 203446
‹treatment-parameter› | MDC_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Setting, ‹CheckingNecessary› ¦ FunctionalDisturbance ¦ Device | Advisory | Setting, ‹CheckingNecessary› | FunctionalDisturbance | Device | tech |  | The user should check an unspecified treatment setting. | O | 0 |  | EVT(0) | 3 | 6838 | 3::6838 | 203446
num_pd_volume_UF_total | MDC_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | tech |  | Total UF volume is too low. | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
num_pd_drain_volume_initial_setting | MDC_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | tech |  | Initial drain volume is too low. | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
‹vendor-text› | MDC_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_PDIALY_PRESCRIPTION_ERROR | Invalid ¦ Prescription ¦ System ¦ Dialysis, Peritoneal | Invalid | Prescription | System | Dialysis, Peritoneal | tech |  | Invalid prescription received. | O | 1 | 0 | EVT(0) | 3 | 1806 | 3::1806 | 198414
num_pd_volume_UF_total | MDC_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_PDIALY_ULTRAFILTRATION_NEG | Too Low ¦ UF ¦ Treatment ¦ Dialysis, Peritoneal | Too Low | UF | Treatment | Dialysis, Peritoneal | tech |  | Negative Ultrafiltration | O | 1 | 0 | EVT(0) | 3 | 1808 | 3::1808 | 198416
num_pd_volume_UF_total | MDC_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_PDIALY_ULTRAFILTRATION_POS | Too High ¦ UF ¦ Treatment ¦ Dialysis, Peritoneal | Too High | UF | Treatment | Dialysis, Peritoneal | tech |  | Positive Ultrafiltration | O | 1 | 0 | EVT(0) | 3 | 1810 | 3::1810 | 198418
‹patient-weight› ‹treatment-data› | MDC_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_TREATMENT_DATA_INVALID | Invalid ¦ Treatment Data ¦ System ¦ Device | Invalid | Treatment Data | System | Device | tech |  | Invalid treatment data (e.g. possibly after a power failure) | O | 1 | 0 | EVT(0) | 3 | 1812 | 3::1812 | 198420
num_pd_patient_volume | MDC_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_VOLUME_DEVIATION | Deviation, from expected ¦ Volume ¦ System ¦ Device | Deviation, from expected | Volume | System | Device | tech |  | Patient volume monitoring has identified a deviation | O | 1 | 0 | EVT(0) | 3 | 1814 | 3::1814 | 198422
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_COMM_ERR_EXTERNAL | Error ¦ Communication ¦ External system or device ¦ Device | Error | Communication | External system or device | Device | tech |  | Internal communication error. | O | 1 | 0 | EVT(0) | 3 | 1816 | 3::1816 | 198424
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_COMM_ERR_INTERNAL | Error ¦ Communication ¦ Internal subsystem or component ¦ Device | Error | Communication | Internal subsystem or component | Device | tech |  | Internal communication error. | O | 1 | 0 | EVT(0) | 3 | 1818 | 3::1818 | 198426
‹treatment-data› | MDC_DEV_PDIALY_VMD | MDC_EVT_DISPLAYED_VALUES_INCONSISTENT | Inconsistent values ¦ Display ¦ System ¦ Device | Inconsistent values | Display | System | Device | tech |  | Inconsistent safety-relevant information displayed | O | 1 | 0 | EVT(0) | 3 | 1820 | 3::1820 | 198428
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_DISPOSABLE_SET_LEAK | Leak ¦ Disposable Set ¦ System ¦ Device | Leak | Disposable Set | System | Device | tech |  | Disposable set leak | O | 1 | 0 | EVT(0) | 3 | 1822 | 3::1822 | 198430
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_LOAD_NEW_SET | Advice ¦ Load New Set ¦ System ¦ Device | Advice | Load New Set | System | Device | tech |  | Load a new set | O | 1 | 0 | EVT(0) | 3 | 1824 | 3::1824 | 198432
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_LOAD_NEW_SET_AND_BAGS | Advice ¦ Load New Set And Bags ¦ System ¦ Device | Advice | Load New Set And Bags | System | Device | tech |  | Load new set and bags | O | 1 | 0 | EVT(0) | 3 | 1826 | 3::1826 | 198434
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_ADVIS_TREATMENT_MAY_BE_INSUFFICIENT | Check ¦ Treatment, Insufficient ¦ System ¦ Dialysis, Peritoneal | Check | Treatment, Insufficient | System | Dialysis, Peritoneal | tech |  | Potentially insufficient treatment detected | O | 1 | 4 | EVT(0) | 3 | 6902 | 3::6902 | 203510
‹mds-or-vmd› ‹any-bag-or-line› | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_AIR_DETECTED | Fault ¦ Air Detected ¦ System ¦ Dialysis, Peritoneal | Fault | Air Detected | System | Dialysis, Peritoneal | tech |  | Maximum allowable air detected in cassette or lines. | O | 1 | 0 | EVT(0) | 3 | 1828 | 3::1828 | 198436
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_CASSETTE_ABSENT | Missing ¦ Cassette ¦ System ¦ Dialysis, Peritoneal | Missing | Cassette | System | Dialysis, Peritoneal | tech |  | Cassette not detected | O | 1 | 0 | EVT(0) | 3 | 1830 | 3::1830 | 198438
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_CASSETTE_AIR_DETECTED | Air ¦ Cassette ¦ System ¦ Dialysis, Peritoneal | Air | Cassette | System | Dialysis, Peritoneal | tech |  | Maximum allowable air detected in cassette. | O | 1 | 0 | EVT(0) | 3 | 1832 | 3::1832 | 198440
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_CASSETTE_FAILURE | Fault ¦ Cassette ¦ System ¦ Dialysis, Peritoneal | Fault | Cassette | System | Dialysis, Peritoneal | tech |  | Cassette failure | O | 1 | 0 | EVT(0) | 3 | 1834 | 3::1834 | 198442
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_INITIAL_DRAIN_VOLUME_UNEXPECTED | Deviation ¦ Drain Volume ¦ System ¦ Dialysis, Peritoneal | Deviation | Drain Volume | System | Dialysis, Peritoneal | tech |  | Unexpected volume (of fluid) during initial drain | O | 1 | 0 | EVT(0) | 3 | 1836 | 3::1836 | 198444
‹mds-or-vmd› num_pd_heater_and_drain_tray_weight | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_PRIMING_ERROR | Failure ¦ Priming ¦ System ¦ Dialysis, Peritoneal | Failure | Priming | System | Dialysis, Peritoneal | tech |  | Priming error | O | 1 | 0 | EVT(0) | 3 | 1838 | 3::1838 | 198446
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_RELOAD_SET | Error ¦ Cassette Loading ¦ System ¦ Dialysis, Peritoneal | Error | Cassette Loading | System | Dialysis, Peritoneal | tech |  | Reload the disposable set | O | 1 | 0 | EVT(0) | 3 | 1840 | 3::1840 | 198448
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_TREATMENT_PAUSED_TOO_LONG | Timeout ¦ Paused ¦ System ¦ Dialysis, Peritoneal | Timeout | Paused | System | Dialysis, Peritoneal | tech |  | Treatment delay | O | 1 | 0 | EVT(0) | 3 | 1842 | 3::1842 | 198450
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_PDIALY_TREATMENT_TIME_EXCEEDED | Timeout ¦ Treatment Time ¦ System ¦ Dialysis, Peritoneal | Timeout | Treatment Time | System | Dialysis, Peritoneal | tech |  | Treatment time exceeded | O | 1 | 0 | EVT(0) | 3 | 1844 | 3::1844 | 198452
‹mds› | MDC_DEV_PDIALY_VMD | MDC_EVT_PT_SECURITY_KEY_ABSENT | Missing ¦ Security Key, Patient Data ¦ System ¦ Device | Missing | Security Key, Patient Data | System | Device | tech |  | Patient security key or other credential is missing. | O | 1 | 0 | EVT(0) | 3 | 1846 | 3::1846 | 198454
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_STAT_DOOR_NOT_CLOSED_OR_LOCKED_CORRECTLY | Invalid ¦ Door Position ¦ System ¦ Device | Invalid | Door Position | System | Device | tech |  | Door is not closed or locked correctly. | O | 1 | 2 | EVT(0) | 3 | 6434 | 3::6434 | 203042
‹mds› | MDC_DEV_PDIALY_VMD | MDC_EVT_STAT_PDIALY_HEATER_BAG_FILLING | Status ¦ Fill, Heater Bag ¦ System ¦ Dialysis, Peritoneal | Status | Fill, Heater Bag | System | Dialysis, Peritoneal | tech |  | Filling heater bag | O | 1 | 2 | EVT(0) | 3 | 6436 | 3::6436 | 203044
‹mds-or-vmd-or-chan› | MDC_DEV_PDIALY_VMD | MDC_EVT_STAT_PDIALY_PT_DISCONN | Status ¦ Disconnected ¦ System ¦ Dialysis, Peritoneal | Status | Disconnected | System | Dialysis, Peritoneal | tech |  | Patient disconnect timeout (intentional) | O | 1 | 2 | EVT(0) | 3 | 6438 | 3::6438 | 203046
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_STAT_PDIALY_TREATMENT_PAUSED | Status ¦ Paused ¦ System ¦ Dialysis, Peritoneal | Status | Paused | System | Dialysis, Peritoneal | tech |  | Treatment paused (e.g. stop button pressed) | O | 1 | 2 | EVT(0) | 3 | 6440 | 3::6440 | 203048
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_SW_ERR | Error ¦ Software ¦ System ¦ Device | Error | Software | System | Device | tech |  | Software error | O | 1 | 0 | EVT(0) | 3 | 1848 | 3::1848 | 198456
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_TIMER_ERR | Fault ¦ Timer ¦ System ¦ Device | Fault | Timer | System | Device | tech |  | Error during the plausibility check of the time counter | O | 1 | 0 | EVT(0) | 3 | 1850 | 3::1850 | 198458
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_TREATMENT_RESULTS_INCONSISTENT | Inconsistent ¦ Treatment Results ¦ System ¦ Device | Inconsistent | Treatment Results | System | Device | tech |  | Inconsistent treatment results | O | 1 | 0 | EVT(0) | 3 | 1852 | 3::1852 | 198460
‹mds-or-vmd› | MDC_DEV_PDIALY_VMD | MDC_EVT_UNRECOV_TREATMENT_DATA_ERR | Unrecoverable ¦ Treatment, Data ¦ System ¦ Device | Unrecoverable | Treatment, Data | System | Device | tech |  | An unrecoverable data error occurred during treatment | O | 1 | 0 | EVT(0) | 3 | 1854 | 3::1854 | 198462
num_pd_drain_volume | MDC_PDIALY_DRAIN_VOLUME | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | tech |  | Drain volume too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
num_pd_drain_volume | MDC_PDIALY_DRAIN_VOLUME | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | tech |  | Drain volume too low | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
num_pd_volume_drain_phase | MDC_PDIALY_DRAIN_VOLUME | MDC_EVT_PDIALY_DRAIN_PHASE_NOT_COMPLETED | Incomplete ¦ Drain ¦ System ¦ Dialysis, Peritoneal | Incomplete | Drain | System | Dialysis, Peritoneal | tech |  | Drain phase not completed | O | 1 | 0 | EVT(0) | 3 | 1856 | 3::1856 | 198464
num_pd_volume_day_fill | MDC_PDIALY_FILL_VOLUME_SETTING | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Setting, ‹CheckingNecessary› ¦ FunctionalDisturbance ¦ Device | Advisory | Setting, ‹CheckingNecessary› | FunctionalDisturbance | Device | tech |  | User should check the day fill volume setting. | O | 0 |  | EVT(0) | 3 | 6838 | 3::6838 | 203446
num_pd_volume_fill_except_last | MDC_PDIALY_FILL_VOLUME_SETTING | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Setting, ‹CheckingNecessary› ¦ FunctionalDisturbance ¦ Device | Advisory | Setting, ‹CheckingNecessary› | FunctionalDisturbance | Device | tech |  | User should check the fill volume setting. | O | 0 |  | EVT(0) | 3 | 6838 | 3::6838 | 203446
num_pd_volume_last_fill | MDC_PDIALY_FILL_VOLUME_SETTING | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Setting, ‹CheckingNecessary› ¦ FunctionalDisturbance ¦ Device | Advisory | Setting, ‹CheckingNecessary› | FunctionalDisturbance | Device | tech |  | User should check the last fill volume setting. | O | 0 |  | EVT(0) | 3 | 6838 | 3::6838 | 203446
num_pd_volume_night_fill | MDC_PDIALY_FILL_VOLUME_SETTING | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Setting, ‹CheckingNecessary› ¦ FunctionalDisturbance ¦ Device | Advisory | Setting, ‹CheckingNecessary› | FunctionalDisturbance | Device | tech |  | User should check the nigh fill volume setting. | O | 0 |  | EVT(0) | 3 | 6838 | 3::6838 | 203446
num_pd_volume_fill_phase num_pd_volume_refill_phase | MDC_PDIALY_FILL_VOLUME | MDC_EVT_PDIALY_FILL_PHASE_NOT_COMPLETED | Incomplete ¦ Fill ¦ System ¦ Dialysis, Peritoneal | Incomplete | Fill | System | Dialysis, Peritoneal | tech |  | Fill or refill phase not completed | O | 1 | 0 | EVT(0) | 3 | 1760 | 3::1760 | 198368
| MDC_PDIALY_FLUID_LEAK_SENSOR_STATE | MDC_EVT_LEAK | ErrorEvent ¦ ‹Leakage› ¦ FunctionalDisturbance ¦ Device | ErrorEvent | ‹Leakage› | FunctionalDisturbance | Device | tech |  | Leak sensor has detected a leak. | O | 0 |  | EVT(0) | 3 | 60 | 3::60 | 196668
| MDC_PDIALY_FLUID_LEAK_SENSOR_STATE | MDC_EVT_SELFTEST_FAILURE | ErrorEvent ¦ Selftest, ‹Failure› ¦ Functional Disturbance ¦ Device | ErrorEvent | Selftest, ‹Failure› | Functional Disturbance | Device | tech |  | Leak sensor failed selftest. | O | 0 |  | EVT(0) | 3 | 1608 | 3::1608 | 198216
num_pd_drain_pressure | MDC_PDIALY_PRESSURE_DRAIN | MDC_EVT_MSMT_OUT_OF_RANGE | ErrorEvent ¦ Measurement, ‹Out of range› ¦ Processing ¦ Device | ErrorEvent | Measurement, ‹Out of range› | Processing | Device | tech |  | Drain pressure out of range | O | 0 |  | EVT(0) | 3 | 1242 | 3::1242 | 197850
num_pd_pressure | MDC_PDIALY_PRESSURE_DRAIN | MDC_EVT_SELFTEST_FAILURE | ErrorEvent ¦ Selftest, ‹Failure› ¦ Functional Disturbance ¦ Device | ErrorEvent | Selftest, ‹Failure› | Functional Disturbance | Device | tech |  | Drain pressure selftest failure | O | 0 |  | EVT(0) | 3 | 1608 | 3::1608 | 198216
num_pd_hydraulic_pressure | MDC_PDIALY_PRESSURE_HYDRAULIC | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | tech |  | Hydraulic pressure too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
num_pd_hydraulic_pressure | MDC_PDIALY_PRESSURE_HYDRAULIC | MDC_EVT_LEAK | ErrorEvent ¦ ‹Leakage› ¦ FunctionalDisturbance ¦ Device | ErrorEvent | ‹Leakage› | FunctionalDisturbance | Device | tech |  | Leakage hydraulic system detected. | O | 0 |  | EVT(0) | 3 | 60 | 3::60 | 196668
num_pd_hydraulic_pressure | MDC_PDIALY_PRESSURE_HYDRAULIC | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | tech |  | Hydraulic pressure too low | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
num_pd_hydraulic_pressure | MDC_PDIALY_PRESSURE_HYDRAULIC | MDC_EVT_SELFTEST_FAILURE | ErrorEvent ¦ Selftest, ‹Failure› ¦ Functional Disturbance ¦ Device | ErrorEvent | Selftest, ‹Failure› | Functional Disturbance | Device | tech |  | Hydraulic pressure selftest failure | O | 0 |  | EVT(0) | 3 | 1608 | 3::1608 | 198216
num_pd_patient_pressure | MDC_PDIALY_PRESSURE_PATIENT | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | tech |  | Patient pressure too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
num_pd_patient_pressure | MDC_PDIALY_PRESSURE_PATIENT | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | tech |  | Patient Pressure too low | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
num_pd_patient_pressure | MDC_PDIALY_PRESSURE_PATIENT | MDC_EVT_MSMT_ERR | ErrorEvent ¦ Measurement, ‹Error› ¦ Processing ¦ Device | ErrorEvent | Measurement, ‹Error› | Processing | Device | tech |  | Unspecified error in patient pressure | O | 0 |  | EVT(0) | 3 | 354 | 3::354 | 196962
‹pump-pressure-sensor› | MDC_PDIALY_PRESSURE_PATIENT | MDC_EVT_SELFTEST_FAILURE | ErrorEvent ¦ Selftest, ‹Failure› ¦ Functional Disturbance ¦ Device | ErrorEvent | Selftest, ‹Failure› | Functional Disturbance | Device | tech |  | Patient pressure selftest failure | O | 0 |  | EVT(0) | 3 | 1608 | 3::1608 | 198216
num_pd_patient_pressure | MDC_PDIALY_PRESSURE_PATIENT | MDC_EVT_HDIALY_PT_LINE_PRESS_HI_LOW_POSN_CHK_REQD | Out of Range ¦ Patient Line Pressure ¦ System ¦ Dialysis, Peritoneal | Out of Range | Patient Line Pressure | System | Dialysis, Peritoneal | tech |  | Patient line pressure too high or low; check patient position | O | 1 | 0 | EVT(0) | 3 | 1858 | 3::1858 | 198466
num_pd_pneumatic_pressure | MDC_PDIALY_PRESSURE_PNEUMATIC | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | tech |  | Pneumatic pressure too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
num_pd_pneumatic_pressure | MDC_PDIALY_PRESSURE_PNEUMATIC | MDC_EVT_LEAK | ErrorEvent ¦ ‹Leakage› ¦ FunctionalDisturbance ¦ Device | ErrorEvent | ‹Leakage› | FunctionalDisturbance | Device | tech |  | Leakage in a gas or fluid filled system detected. | O | 0 |  | EVT(0) | 3 | 60 | 3::60 | 196668
num_pd_pneumatic_pressure | MDC_PDIALY_PRESSURE_PNEUMATIC | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | tech |  | Pneumatic pressure too low | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
num_pd_pressure | MDC_PDIALY_PRESSURE_PNEUMATIC | MDC_EVT_SELFTEST_FAILURE | ErrorEvent ¦ Selftest, ‹Failure› ¦ Functional Disturbance ¦ Device | ErrorEvent | Selftest, ‹Failure› | Functional Disturbance | Device | tech |  | Pneumatic pressure selftest failure | O | 0 |  | EVT(0) | 3 | 1608 | 3::1608 | 198216
num_pd_pneumatic_pressure | MDC_PDIALY_PRESSURE_PNEUMATIC | MDC_EVT_PDIALY_VALVE_LEAK | Leak ¦ Valve ¦ System ¦ Dialysis, Peritoneal | Leak | Valve | System | Dialysis, Peritoneal | tech |  | Valve or pressure leak | O | 1 | 0 | EVT(0) | 3 | 1860 | 3::1860 | 198468
num_pd_vacuum_pressure | MDC_PDIALY_PRESSURE_VACUUM | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | tech |  | Vacuum pressure too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
num_pd_vacuum_pressure | MDC_PDIALY_PRESSURE_VACUUM | MDC_EVT_LEAK | ErrorEvent ¦ ‹Leakage› ¦ FunctionalDisturbance ¦ Device | ErrorEvent | ‹Leakage› | FunctionalDisturbance | Device | tech |  | Leakage in a gas or fluid filled system detected. | O | 0 |  | EVT(0) | 3 | 60 | 3::60 | 196668
num_pd_vacuum_pressure | MDC_PDIALY_PRESSURE_VACUUM | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | tech |  | Vacuum pressure too low | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
num_pd_vacuum_pressure | MDC_PDIALY_PRESSURE_VACUUM | MDC_EVT_MSMT_OUT_OF_RANGE | ErrorEvent ¦ Measurement, ‹Out of range› ¦ Processing ¦ Device | ErrorEvent | Measurement, ‹Out of range› | Processing | Device | tech |  | Vacuum pressure out of range | O | 0 |  | EVT(0) | 3 | 1242 | 3::1242 | 197850
num_pd_pressure | MDC_PDIALY_PRESSURE_VACUUM | MDC_EVT_SELFTEST_FAILURE | ErrorEvent ¦ Selftest, ‹Failure› ¦ Functional Disturbance ¦ Device | ErrorEvent | Selftest, ‹Failure› | Functional Disturbance | Device | tech |  | Vacuum pressure selftest failure | O | 0 |  | EVT(0) | 3 | 1608 | 3::1608 | 198216
| MDC_PDIALY_TEMP_DIALYSATE_IN | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | tech |  | Dialysate inlet temperature too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
| MDC_PDIALY_TEMP_DIALYSATE_IN | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | tech |  | Dialysate inlet temperature too low | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
num_pd_dialysate_temperature | MDC_PDIALY_TEMP_DIALYSATE_OUT | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | tech |  | Dialysate temperature too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
num_pd_dialysate_temperature | MDC_PDIALY_TEMP_DIALYSATE_OUT | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | tech |  | Dialysate temperature too low | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
num_pd_dialysate_temperature | MDC_PDIALY_TEMP_DIALYSATE_OUT | MDC_EVT_MSMT_ERR | ErrorEvent ¦ Measurement, ‹Error› ¦ Processing ¦ Device | ErrorEvent | Measurement, ‹Error› | Processing | Device | tech |  | Unspecified error in dialysate temperature | O | 0 |  | EVT(0) | 3 | 354 | 3::354 | 196962
num_pd_dialysate_temperature | MDC_PDIALY_TEMP_DIALYSATE_OUT | MDC_EVT_MSMT_OUT_OF_RANGE | ErrorEvent ¦ Measurement, ‹Out of range› ¦ Processing ¦ Device | ErrorEvent | Measurement, ‹Out of range› | Processing | Device | tech |  | Dialysate temperature out of range | O | 0 |  | EVT(0) | 3 | 1242 | 3::1242 | 197850
num_pd_dialysate_temperature | MDC_PDIALY_TEMP_DIALYSATE_OUT | MDC_EVT_STAT_SOLUTION_WARMING | Status ¦ Warming, Fluid ¦ Temperature may be too low ¦ Device | Status | Warming, Fluid | Temperature may be too low | Device | tech |  | Dialysate solution is warming | O | 1 | 2 | EVT(0) | 3 | 6442 | 3::6442 | 203050
num_pd_heater_temperature | MDC_PDIALY_TEMP_HEATER_BAG | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | tech |  | Heater bag temperature too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
num_pd_heater_temperature | MDC_PDIALY_TEMP_HEATER_BAG | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | tech |  | Heater bag temperature too low | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
num_pd_temperature | MDC_PDIALY_TEMP_HEATER_BAG | MDC_EVT_MSMT_OUT_OF_RANGE | ErrorEvent ¦ Measurement, ‹Out of range› ¦ Processing ¦ Device | ErrorEvent | Measurement, ‹Out of range› | Processing | Device | tech |  | Heater Bag Temperature out of range | O | 0 |  | EVT(0) | 3 | 1242 | 3::1242 | 197850
num_pd_heater_temperature | MDC_PDIALY_TEMP_HEATER_BAG | MDC_EVT_MSMT_OUT_OF_RANGE_HI | ErrorEvent ¦ Measurement, ‹Out of range›, ‹High› ¦ SignalQuality ¦ Device | ErrorEvent | Measurement, ‹Out of range›, ‹High› | SignalQuality | Device | tech |  | Heater bag temperature too high | O | 0 |  | EVT(0) | 3 | 1244 | 3::1244 | 197852
num_pd_heater_temperature | MDC_PDIALY_TEMP_HEATER_BAG | MDC_EVT_SELFTEST_FAILURE | ErrorEvent ¦ Selftest, ‹Failure› ¦ Functional Disturbance ¦ Device | ErrorEvent | Selftest, ‹Failure› | Functional Disturbance | Device | tech |  | Heater bag temperature selftest failure | O | 0 |  | EVT(0) | 3 | 1608 | 3::1608 | 198216
num_pd_heater_temperature | MDC_PDIALY_TEMP_HEATER_BAG | MDC_EVT_TEMP_ERR | Error ¦ Temperature ¦ System ¦ Device | Error | Temperature | System | Device | tech |  | Temperature sensor error | O | 1 | 0 | EVT(0) | 3 | 1862 | 3::1862 | 198470
num_pd_volume_tidal_setting | MDC_PDIALY_TIDAL_PERCENTAGE | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Setting, ‹CheckingNecessary› ¦ FunctionalDisturbance ¦ Device | Advisory | Setting, ‹CheckingNecessary› | FunctionalDisturbance | Device | tech |  | The user should check the tidal setting. | O | 0 |  | EVT(0) | 3 | 6838 | 3::6838 | 203446
‹any-valve› | MDC_DEV_PDIALY_VMD | MDC_EVT_VALVE_MALF | Malfunction ¦ Valve ¦ System ¦ Device | Malfunction | Valve | System | Device | tech |  | Valve malfunction | O | 1 | 0 | EVT(0) | 3 | 1864 | 3::1864 | 198472
num_pd_heater_and_drain_tray_weight | MDC_PDIALY_WEIGHT_HEATER_BAG | MDC_EVT_HI_GT_LIM | LimitEvent ¦ ‹High, val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High, val>lim› | Processing | Device | tech |  | A metric exceeds a given threshold | O | 0 |  | EVT(0) | 3 | 42 | 3::42 | 196650
num_pd_heater_and_drain_tray_weight | MDC_PDIALY_WEIGHT_HEATER_BAG | MDC_EVT_MSMT_ERR | ErrorEvent ¦ Measurement, ‹Error› ¦ Processing ¦ Device | ErrorEvent | Measurement, ‹Error› | Processing | Device | tech |  | Unspecified error in heater bag weight | O | 0 |  | EVT(0) | 3 | 354 | 3::354 | 196962
num_pd_heater_and_drain_tray_weight | MDC_PDIALY_WEIGHT_HEATER_BAG | MDC_EVT_MSMT_FAIL | ErrorEvent ¦ Measurement, ‹Failed› ¦ FunctionalDisturbance ¦ Device | ErrorEvent | Measurement, ‹Failed› | FunctionalDisturbance | Device | tech |  | Failure in heater bag weight | O | 0 |  | EVT(0) | 3 | 356 | 3::356 | 196964
num_pd_heater_and_drain_tray_weight | MDC_PDIALY_WEIGHT_HEATER_BAG | MDC_EVT_PDIALY_ABRUPT_WEIGHT_CHANGE | Abrupt Change ¦ Weight ¦ System ¦ Dialysis, Peritoneal | Abrupt Change | Weight | System | Dialysis, Peritoneal | tech |  | Abrupt weight change; possible scale interference. | O | 1 | 0 | EVT(0) | 3 | 1866 | 3::1866 | 198474
num_pd_heater_bag_weight num_pd_heater_and_drain_tray_weight | MDC_PDIALY_WEIGHT_HEATER_BAG | MDC_EVT_PDIALY_FLUID_BAG_WEIGHT_VOLUME_ERROR | Inconsistent ¦ Heater Bag Weight ¦ System ¦ Dialysis, Peritoneal | Inconsistent | Heater Bag Weight | System | Dialysis, Peritoneal | tech |  | Fluid bag weight scale (volume) measurement error. | O | 1 | 0 | EVT(0) | 3 | 1868 | 3::1868 | 198476
num_pd_heater_bag_weight | MDC_PDIALY_WEIGHT_HEATER_BAG | MDC_EVT_PDIALY_HEATER_BAG_ON_HEATER_TRAY_UNEXPECTED | Deviation ¦ Heater Bag ¦ System ¦ Dialysis, Peritoneal | Deviation | Heater Bag | System | Dialysis, Peritoneal | tech |  | Heater bag detected on the heater tray when not expected. | O | 1 | 0 | EVT(0) | 3 | 1870 | 3::1870 | 198478
num_pd_heater_and_drain_tray_weight | MDC_PDIALY_WEIGHT_HEATER_BAG | MDC_EVT_WEIGHT_INVALID_CHK_REQD | Weight ¦ Measurement, ‹Invalid›, CheckingRequired ¦ System ¦ Device | Weight | Measurement, ‹Invalid›, CheckingRequired | System | Device | tech |  | The dialysis machine has detected an invalid weight change and the user must verify/confirm the weight. | O | 1 | 0 | EVT(0) | 3 | 1872 | 3::1872 | 198480
| MDC_PRESS_BLD_NONINV_DIA | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | phys | Diastolic High | Diastolic pressure is too high. | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
| MDC_PRESS_BLD_NONINV_DIA | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | phys | Diastolic Low | Diastolic pressure is too low. | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
| MDC_PRESS_BLD_NONINV_SYS | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | phys | Systolic High | Systolic pressure is too high. | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
| MDC_PRESS_BLD_NONINV_SYS | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | phys | Systolic Low | Systolic pressure is too low. | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
| MDC_PULS_RATE_NON_INV | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | phys | Pulse High | Pulse is too high. | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
| MDC_PULS_RATE_NON_INV | MDC_EVT_LO | LimitEvent ¦ ‹Low›, ‹val<lim› ¦ Processing ¦ Device | LimitEvent | ‹Low›, ‹val<lim› | Processing | Device | phys | Pulse Low | Pulse is too low. | O | 0 |  | EVT(0) | 3 | 62 | 3::62 | 196670
| MDC_TEMP_TYMP | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | phys | Temperature High | Temperature is too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
| MDC_TEMP_ORAL | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | phys | Temperature High | Temperature is too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
| MDC_TEMP_BODY | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | phys | Temperature High | Temperature is too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648
| MDC_TEMP_AXILLA | MDC_EVT_HI | LimitEvent ¦ ‹High›, ‹val>lim› ¦ Processing ¦ Device | LimitEvent | ‹High›, ‹val>lim› | Processing | Device | phys | Temperature High | Temperature is too high | O | 0 |  | EVT(0) | 3 | 40 | 3::40 | 196648

## ConditionCodes
Code | Condition | Expression
C1 | The Exchange Phase is Fill or greater | MDCX_PDIALY_CURRENT_PHASE = FILL or DWELL or DRAIN or COMPLETE
C2 | The Exchange Phase is Dwell or greater | MDCX_PDIALY_CURRENT_PHASE = DWELL or DRAIN or COMPLETE
C3 | The Exchange Phase is Drain or greater | MDCX_PDIALY_CURRENT_PHASE = DRAIN or COMPLETE
C4 | The Fluid Source is bags. | MDCX_PDIALY_FLUID_SOURCE = BAG
C5 | There are APD exchanges | MDCX_PDIALY_TREAT_TYPE = APD
C6 | There are CAPD exchanges | MDCX_PDIALY_TREAT_TYPE = CAPD
C7 | The prescription type is exchange based. | MDCX_PDIALY_PRESCRIPTION_TYPE = EXCH
C11 | APD Fixed Duration Prescription, (MDC_PDIALY_PRESCRIPTION_TYPE = TXDUR) & (MDC_PDIALY_TREAT_TYPE = APD).
C12 | APD Treatment Based Prescription, ((MDC_PDIALY_PRESCRIPTION_TYPE = TXDUR) ¦ (MDC_PDIALY_PRESCIPTION_TYPE = TXEND)) & (MDC_PDIALY_TREAT_TYPE = CPD)
C13 | CPD Fixed Duration Prescription, (MDC_PDIALY_PRESCRIPTION_TYPE = TXDUR) & (MDC_PDIALY_TREAT_TYPE = APD).
C14 | CPD Treatment Based Prescription, ((MDC_PDIALY_PRESCRIPTION_TYPE = TXDUR) ¦ (MDC_PDIALY_PRESCIPTION_TYPE = TXEND)) & (MDC_PDIALY_TREAT_TYPE = CPD)