## Data
REFID | SysName | SysNameBase | SysName1stDesc | SysName2ndDesc | SysName3rdDesc | AlertType | CommonTerm | Description | Phase | Temporal | DataType | Format | UOM | UOM_UCUM | Card | Use | RxUse | Enum_Table | Enum_Values | Enum_Descriptions | Range | Example | gp | s | Disc | PART | pCODE10 | CODE10 | CF_CODE10
MDCX_DEV_PDIALY_MACHINE_MDS | System ¦ ¦ ¦ Dialysis, Peritoneal | System |  |  | Dialysis, Peritoneal |  | Hemodialysis machine | A device or system used to treat renal failure delivered by removing toxins from the blood through an artificial kidney, called a dialyzer. | All | All |  |  |  |  | 1..1 | M | M
. . . MDC_ID_MODEL_MANUFACTURER | System (string) ¦ Identifier ¦ Manufacturer ¦ Device | System (string) | Identifier | Manufacturer | Device |  | Dialysis Device Manufacturer | String containing device manufacturer's name | All | All | String | Alphanumeric | N/A | -- | 1..1 | M | X | -- |  | -- | N/A | NxStage
. . . MDC_ID_MODEL_NUMBER | System (string) ¦ Identifier ¦ Model Number ¦ Device | System (string) | Identifier | Model Number | Device |  | Dialysis Device Model | String containing device manufacturer's model identifier for the device | All | All | String | Alphanumeric | N/A | -- | 1..1 | M | X | -- |  | -- | N/A | System One
. . . MDC_ID_PROD_SPEC_SERIAL | System (string) ¦ Identifier ¦ Serial Number ¦ Device | System (string) | Identifier | Serial Number | Device |  | Dialysis Device Serial Number | String containing the device’s serial number | All | All | String | Alphanumeric | N/A | -- | 1..1 | M | X | -- |  | -- | N/A | AS1000478
. . . MDC_ID_PROD_SPEC_SW | System (string) ¦ Identifier ¦ Software Version ¦ Device | System (string) | Identifier | Software Version | Device |  | Device Software Version | String defining the version of the software on the machine | All | All | String | Alphanumeric | N/A | -- | 1..1 | M | X | -- |  | -- | N/A | 1.2.3.4
. . . MDC_ATTR_ID_UDI | System (string) ¦ Identifier, Unique ¦ Universal identifier, UDI ¦ Device | System (string) | Identifier, Unique | Universal identifier, UDI | Device |  | Dialysis Device UDI | String containing the device’s FDA required Universal Device Identifier . [Note: MDC_ATTR_ID_UDI contains three components: UdiAuthority, UdiIssuer, UdiLabel to support non-FDA entities, but currently there is no approved way of encoding this using HL7 V2; PRT-10 and PRT-16-20 is used instead; UdiLabel = PRT-10.] | All | All | String | Alphanumeric | N/A | -- | 1..1 | O | X | -- |  | -- | N/A | +M535NX10003A0/$$+735241/16D20180305J
. MDCX_DEV_PDIALY_VMD | Device ¦ ¦ ¦ Dialysis, Peritoneal | Device |  |  | Dialysis, Peritoneal |  | Hemodialysis subsystem | A device or subsystem used to treat renal failure delivered by removing toxins from the blood through an artificial kidney, called a dialyzer. | All | All |  |  |  |  | 1..1 | M | M
. . MDCX_DEV_PDIALY_MACH_CONFIG_CHAN | Subsystem ¦ ¦ Machine Configuration ¦ Dialysis, Peritoneal | Subsystem |  | Machine Configuration | Dialysis, Peritoneal |  | Hemodialysis configuration | The principal dialysis machine operational modes, treatment modality, settings and other information about the dialysis machine. | All | All |  |  |  |  | 1..1 | M | X
. . . MDC_ATTR_TIME_ABS | Time ¦ Machine Time ¦ Machine Configuration ¦ Dialysis | Time | Machine Time | Machine Configuration | Dialysis |  | Dialysis Device Time | Date and time as recorded on the dialysis device's internal clock with offset to UTC (based on location and DST) | All | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 1..1 | M | X | -- |  | -- | N/A | 20190421130045+/-0000 |  |  |  | 1 | 2439 | 67975
. . MDCX_DEV_PDIALY_TREATMENT_CHAN | Subsystem ¦ Treatment ¦ ¦ Dialysis, Peritoneal | Subsystem | Treatment |  | Dialysis, Peritoneal |  | Treatment |  | Intradialytic | All |  |  |  |  | 1..1 | M | M
. . . MDCX_PDIALY_TREAT_TYPE | Treatment Type ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Treatment Type | Setting | Treatment | Dialysis, Peritoneal |  | Treatment Type | Treatment Type | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 1..1 | M | M | PD_TBL_01 | CAPD APD CCPD NIPD PDplus TPD aAPD IPD Other | CAPD - Continuous Ambulatory Peritoneal Dialysis APD _ Automated Peritoneal Dialysis CCPD - Continuous Cycling Peritoneal Dialysis NIPD - Nocturnal Intermittent Peritoneal Dialysis PDPLUS - PD Plus TPD - Tidal Peritoneal Dialysis AAPD - Adapted Automated Peritoneal Dialysis IPD - Intermittent Peritoneal Dialysis OTHER - Undefined | N/A | CAPD
. . . MDCX_PDIALY_TREAT_LOCATION | Treatment Location ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | Treatment Location | Status | Treatment | Dialysis, Peritoneal |  | Treatment Location | Treatment Location | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 0..1 | O | O | PD_TBL_02 | HOME CLINIC OTHER | Home Clinic Other/Unknown | N/A | HOME
. . . MDCX_PDIALY_CURRENT_PHASE | Current Phase ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | Current Phase | Status | Treatment | Dialysis, Peritoneal |  | Current Phase | The current treatment phase. | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 1..1 | M | X | PD_TBL_03 | PENDING FILL DWELL DRAIN COMPLETE | Pending Fill Dwell Drain Complete | N/A | COMPLETE
. . . MDCX_PDIALY_PRESCRIPTION_TYPE | Type ¦ Status ¦ Prescription ¦ Dialysis, Peritoneal | Type | Status | Prescription | Dialysis, Peritoneal |  | Prescription Type | Prescription Type | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 1..1 | M | M | PD_TBL_13 | EXCH TXDUR TXEND | Exchange Based Treatment Based - Fixed Duration Treatment Based - Fixed End Time | N/A | EXCH
. . . MDCX_PDIALY_APD_START_TIME_SETTING | ADP Start Time ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | ADP Start Time | Status | Treatment | Dialysis, Peritoneal |  | APD Start Time | APD Start Time | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 1..1 | C5 | C5 | N/A |  |  |  | 20190421130045+/-0000
. . . MDCX_PDIALY_APD_END_TIME_SETTING | ADP End Time ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | ADP End Time | Status | Treatment | Dialysis, Peritoneal |  | APD End Time | APD End Time | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 1..1 | C5 | C5 | N/A |  |  |  | 20190421130045+/-0000
. . . MDCX_PDIALY_APD_TOTAL_VOLUME_SETTING | ADP Volume ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | ADP Volume | Status | Treatment | Dialysis, Peritoneal |  | APD Total Volume | Total volume in APD exchanges. | Intradialytic | All | Numeric | X.XXX | L | L | 1..1 | C5 | C5 | N/A
. . . MDCX_PDIALY_APD_CYCLES_SETTING | ADP Cycler ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | ADP Cycler | Setting | Treatment | Dialysis, Peritoneal |  | APD Cycles | APD Cycles | Intradialytic | All | Numeric | XX | N/A | -- | 1..1 | C5 | C5 | N/A
. . . MDCX_PDIALY_CAPD_START_TIME_SETTING | CAPD Start Time ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | CAPD Start Time | Status | Treatment | Dialysis, Peritoneal |  | CAPD Start Time | CAPD Start Time | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 1..1 | C6 | C6 | N/A |  |  |  | 20190421130045+/-0000
. . . MDCX_PDIALY_CAPD_END_TIME_SETTING | CAPD End Time ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | CAPD End Time | Status | Treatment | Dialysis, Peritoneal |  | CAPD End Time | CAPD End Time | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 1..1 | C6 | C6 | N/A |  |  |  | 20190421130045+/-0000
. . . MDCX_PDIALY_CAPD_TOTAL_VOLUME_SETTING | CAPD Volume ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | CAPD Volume | Status | Treatment | Dialysis, Peritoneal |  | CAPD Exchanges | CAPD Exchanges | Intradialytic | All | Numeric |  | N/A | -- | 1..1 | C6 | C6 | N/A
. . . MDCX_PDIALY_CAPD_CYCLES_SETTING | CAPD Cycles ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | CAPD Cycles | Setting | Treatment | Dialysis, Peritoneal |  | CAPD Total Volume | Total volume in CAPD exchanges | Intradialytic | All | Numeric | XX | N/A | -- | 1..1 | C6 | C6 | N/A
. . . MDCX_PDIALY_TIDAL_PERCENTAGE | Tidal Percentage ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Tidal Percentage | Setting | Treatment | Dialysis, Peritoneal |  | Tidal Percentage | The percentage of Night Fill Volume to be drained during a tidal cycle. | Intradialytic | All | Numeric | XXX.X | % | % | 0..1 | O | O | N/A
. . . MDCX_PDIALY_FULL_DRAIN_FREQ | Full Drain Frequency ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Full Drain Frequency | Setting | Treatment | Dialysis, Peritoneal |  | Full Drain Frequency | The number of cycles between full drains when performing tidal exchanges. The values is expressed as the number of drains between full drains. | Intradialytic | All | Numeric | X | N/A | -- | 0..1 | O | O | N/A
. . . MDCX_PDIALY_PERITONEUM_VOL_MAX | Peritoneum Max Volume ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Peritoneum Max Volume | Setting | Treatment | Dialysis, Peritoneal |  | Maximum Peritoneum Volume |  | Intradialytic | All | Numeric | X.XXX | L | L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_INIT_MIN_DRAIN_VOL_SETTING | Drain Min Volume ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Drain Min Volume | Setting | Treatment | Dialysis, Peritoneal |  | Minimum Initial Drain Volume | The minimum volume to remove during the initial drain. | Intradialytic | All | Numeric | X.XXX | L | L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_INIT_MAX_DRAIN_VOL_SETTING | Drain Max Volume ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Drain Max Volume | Setting | Treatment | Dialysis, Peritoneal |  | Maximum Initial Drain Time | The maximum volume to remove during the initial drain. | Intradialytic | All | Numeric | X.XXX | L | L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_SAMPLE_REMINDER | Sample Reminder Required ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Sample Reminder Required | Setting | Treatment | Dialysis, Peritoneal |  | Sample Reminder | This value is true when the user should be reminded to collect an effluent sample. | Intradialytic | All | Bool | T / F | N/A | -- | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_MEDICATION_SURVEY | Medication Survey Required ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Medication Survey Required | Setting | Treatment | Dialysis, Peritoneal |  | Medications Survey | This value is true when the user should be surveyed for medications taken. | Intradialytic | All | Bool | T / F | N/A | -- | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_EXCHANGE_SURVEY | Exchange Survey Required ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Exchange Survey Required | Setting | Treatment | Dialysis, Peritoneal |  | Exchange Survey | This value is true when the user should be surveyed for number of day exchanges performed. | Intradialytic | All | Bool | T / F | N/A | -- | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_DAY_DRAIN_TIME_MIN | Minimum Drain Time ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Minimum Drain Time | Setting | Treatment | Dialysis, Peritoneal |  | Minimum Day Drain Time | Minimum duration for a day time drain. | Intradialytic | All | Numeric | XXX.X | min | min | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_DAY_DRAIN_VOL_PCT_MIN | Minimum Drain Volume Percentage ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Minimum Drain Volume Percentage | Setting | Treatment | Dialysis, Peritoneal |  | Minimum Day Drain Volume Percent | The min amount of fluid to drain expressed as a percentage of the fill volume. | Intradialytic | All | Numeric | XXX.X | % |  | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_NIGHT_DRAIN_TIME_MIN | Minimum Night Drain Time ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Minimum Night Drain Time | Setting | Treatment | Dialysis, Peritoneal |  | Minimum Night Drain Time | Minimum duration for a night time drain. | Intradialytic | All | Numeric | XXX.X | min | min | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_LAST_DRAIN_MODE | Last Drain Mode ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Last Drain Mode | Setting | Treatment | Dialysis, Peritoneal |  | Extra Last Drain Mode | This value specifies how the last drain phase is performed. | Intradialytic | All | Enumeration |  | N/A | -- | 0..1 | O | O | PD_TBL_12 | NEVER ALWAYS UFLIMIT | Never - It will never take place Always - It will always take place UF Limit - It will take place if the total estimated UF for the therapy is below the Extra Drain UF Limit. The Extra Drain UF Limit is expressed as a percentage of total estimated UF for the therapy.
. . . MDCX_PDIALY_LAST_DRAIN_UF_LIMIT | Last Drain UF Limit ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Last Drain UF Limit | Setting | Treatment | Dialysis, Peritoneal |  | Extra Last Drain UF Limit | This limit which determines if the extra drain is performed. If the total estimated UF for the therapy is below the Extra Drain UF Limit the extra drain is performed. The Extra Drain UF Limit is expressed as a percentage of total estimated UF for the therapy. | Intradialytic | All | \ | XXX.X | % | % | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_NIGHT_UF_ESTIMATE | Night UF Estimate ¦ Status ¦ Treatment ¦ Dialysis, Peritoneal | Night UF Estimate | Status | Treatment | Dialysis, Peritoneal |  | Estimated Night UF | The estimated UF volume for the night therapy. This value is used to keep track of the fluid balance for the patient, so that the fluid that was filled plus the fluid generated as UF is drained from the patient. | Intradialytic | All | Numeric | X.XXX | L | L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_TEMP_DIALYSATE_SETTING | Dialysate Temperature ¦ Setting ¦ Treatment ¦ Dialysis, Peritoneal | Dialysate Temperature | Setting | Treatment | Dialysis, Peritoneal |  | Dialysate Temperature Setting | Desired dialysate temperature during treatment | Intradialytic | All | Numeric | XX.X | ˚C | Cel | 0..1 | O | X | N/A |  | N/A
. . MDCX_DEV_PDIALY_FLUID_1_CHAN . . MDCX_DEV_PDIALY_FLUID_2_CHAN . . MDCX_DEV_PDIALY_FLUID_3_CHAN . . MDCX_DEV_PDIALY_FLUID_4_CHAN . . MDCX_DEV_PDIALY_FLUID_5_CHAN . . MDCX_DEV_PDIALY_FLUID_6_CHAN | Subsystem ¦ Fluid ¦ ¦ Dialysis, Peritoneal | Subsystem | Fluid |  | Dialysis, Peritoneal |  | Fluid Definition |  | Intradialytic | All |  |  |  |  | 1..1 | M | M
. . . MDCX_PDIALY_FLUID_NAME | Name ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Name | Setting | Fluid | Dialysis, Peritoneal |  | Dialysate Name | Dialysate Name | Intradialytic | All | String | Alphanumeric | N/A | -- | 0..1 | O | O | N/A |  | N/A
. . . MDCX_DIALY_FLUID_GLUCOSE_CONC | Glucose ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Glucose | Setting | Fluid | Dialysis, Peritoneal |  | Glucose | Glucose | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_DIALY_FLUID_ICODEXTRIM_CONC | Icodextrim ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Icodextrim | Setting | Fluid | Dialysis, Peritoneal |  | Icodextrin | Icodextrin | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_DIALY_FLUID_AMINO_ACID_CONC | Amino Acid ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Amino Acid | Setting | Fluid | Dialysis, Peritoneal |  | Amino acid | Amino acid | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_DIALY_FLUID_OSMOTIC_AGENT | Agent ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Agent | Setting | Fluid | Dialysis, Peritoneal |  | Other osmotic agent type | Other osmotic agent type | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 0..1 | O | O | PD_TBL_04 | GLU DEX ICO OTHER | Glucose Dextrose Icodextrin Other
. . . MDCX_DIALY_FLUID_OSMOTIC_AGENT_CONC | Agent Concentration ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Agent Concentration | Setting | Fluid | Dialysis, Peritoneal |  | Other osmotic agent concentration | Other osmotic agent concentration | Intradialytic | All | Numeric | XX.XX | % |  | 0..1 | O | O | N/A |  | N/A
. . . MDCX_DIALY_FLUID_BICARB_CONC | Bicarb ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Bicarb | Setting | Fluid | Dialysis, Peritoneal |  | Bicarbonate (HCO3) | Bicarbonate (HCO3) | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_DIALY_FLUID_CHLORIDE_CONC | Chloride ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Chloride | Setting | Fluid | Dialysis, Peritoneal |  | Chloride (CL) | Chloride (CL) | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_DIALY_FLUID_LACTATE_CONC | Lacatate ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Lacatate | Setting | Fluid | Dialysis, Peritoneal |  | Lactate | Lactate | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_DIALY_FLUID_CALCIUM_CONC | Calcium ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Calcium | Setting | Fluid | Dialysis, Peritoneal |  | Calcium (CA) | Calcium (CA) | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_DIALY_FLUID_CITRATE_CONC | Citrate ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Citrate | Setting | Fluid | Dialysis, Peritoneal |  | Citrate | Citrate | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_DIALY_FLUID_POTASSIUM_CONC | Potassium ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Potassium | Setting | Fluid | Dialysis, Peritoneal |  | Potassium (K) | Potassium (K) | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_DIALY_FLUID_MAGNESIUM_CONC | Magnesium ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Magnesium | Setting | Fluid | Dialysis, Peritoneal |  | Magnesium | Magnesium | Intradialytic | All | Numeric | XXX.X | mmol/L | mmol/L | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_FLUID_BAG_VOLUME | Volume ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Volume | Setting | Fluid | Dialysis, Peritoneal |  | Bag Volume | Bag Volume | Intradialytic | All | Numeric | X.XXX | L | L | 1..1 | C4 | C4 | N/A |  | N/A
. . . MDCX_PDIALY_FLUID_SOURCE | Source ¦ Setting ¦ Fluid ¦ Dialysis, Peritoneal | Source | Setting | Fluid | Dialysis, Peritoneal |  | Fluid Source | Source of the fluid. | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 1..1 | M | M | PD_TBL_05 | BAG ONLINE UNKNOWN | Bag Online Uknown
. . MDCX_DEV_PDIALY_EXCHANGE_CHAN | Subsystem ¦ Exchange ¦ ¦ Dialysis, Peritoneal | Subsystem | Exchange |  | Dialysis, Peritoneal |  | Peritoneal Exchange | The values for the current and previous exchanges. | Intradialytic | All |  |  |  |  | 1..1 | M | C7
. . . MDC_ATTR_CHAN_NUM_LOGICAL | Number ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Number | Setting | Exchange | Dialysis, Peritoneal |  | Exchange Number | The number of the exchange in the current treatment. The first exchange is number 1. | Intradialytic | All | Numeric | X | N/A | -- | 1..1 | M | C7 | -- |  | -- | 1-9 | 1
. . . MDCX_PDIALY_CURRENT_PHASE | Phase ¦ Status ¦ Exchange ¦ Dialysis, Peritoneal | Phase | Status | Exchange | Dialysis, Peritoneal |  | Current Phase | The current exchange phase. | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 1..1 | M | X | PD_TBL_03 |  | See MDCX_PDIALY_CURRENT_PHASE value in the Treatment Channel.
. . . MDCX_PDIALY_FILL_START_TIME | Fill Start ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Fill Start | Time | Exchange | Dialysis, Peritoneal |  | Fill Start time | The time that the fill phase started. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 1..1 | C1 | X | N/A |  | N/A
. . . MDCX_PDIALY_FILL_END_TIME | Fill End ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Fill End | Time | Exchange | Dialysis, Peritoneal |  | Fill End time | The time that the fill phase ended. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 1..1 | C1 | X | N/A |  | N/A
. . . MDCX_PDIALY_FILL_MODE | Fill Mode ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Fill Mode | Setting | Exchange | Dialysis, Peritoneal |  | Fill mode | Fill mode | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 1..1 | C1 | C7 | PD_TBL_06 | AUTO MANUAL | Cycler Manual
. . . MDCX_PDIALY_FILL_DURATION_SETTING | Fill Duration ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Fill Duration | Setting | Exchange | Dialysis, Peritoneal |  | Prescribed Fill Duration | Prescribed Fill duration | Intradialytic | All | Numeric | XXX.X | min | min | 1..1 | C1 | C7 | N/A |  | N/A
. . . MDCX_PDIALY_FILL_DURATION | Fill Duration ¦ Actual ¦ Exchange ¦ Dialysis, Peritoneal | Fill Duration | Actual | Exchange | Dialysis, Peritoneal |  | Actual Fill Duration | Actual fill duration | Intradialytic | All | Numeric | XXX.X | min | min | 1..1 | C1 | X | N/A |  | N/A
. . . MDCX_PDIALY_FILL_VOLUME_SETTING | Fill Volume ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Fill Volume | Setting | Exchange | Dialysis, Peritoneal |  | Prescribed Fill volume | Prescribed Fill volume | Intradialytic | All | Numeric Array | X.XXX | L | L | 1..1 | C1 | C7 | N/A |  | N/A
. . . MDCX_PDIALY_FILL_VOLUME | Fill Volume ¦ Actual ¦ Exchange ¦ Dialysis, Peritoneal | Fill Volume | Actual | Exchange | Dialysis, Peritoneal |  | Actual Fill volume | Actual Fill volume | Intradialytic | All | Numeric Array | X.XXX | L | L | 1..1 | C1 | X | N/A |  | N/A
. . . MDCX_PDIALY_FLUID_TEMP_SETTING | Fluid Temperature ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Fluid Temperature | Setting | Exchange | Dialysis, Peritoneal |  | Fluid temperature setting | Fluid temperature | Intradialytic | All | Numeric | XX.X | degC | Cel | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_FLUID_TEMP | Fluid Temperature ¦ Status ¦ Exchange ¦ Dialysis, Peritoneal | Fluid Temperature | Status | Exchange | Dialysis, Peritoneal |  | Fluid temperature | Fluid temperature | Intradialytic | All | Numeric | XX.X | degC | Cel | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_DWELL_START_TIME | Dwell Start ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Dwell Start | Time | Exchange | Dialysis, Peritoneal |  | Dwell Start time | The time that the dwell phase started. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 1..1 | C2 | X | N/A |  | N/A
. . . MDCX_PDIALY_DWELL_END_TIME | Dwell End ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Dwell End | Time | Exchange | Dialysis, Peritoneal |  | Dwell End time | The time that the dwell phase ended. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 1..1 | C2 | X | N/A |  | N/A
. . . MDCX_PDIALY_DWELL_DURATION_SETTING | Dwell Duration ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Dwell Duration | Setting | Exchange | Dialysis, Peritoneal |  | Prescribed Dwell Duration | Prescribed dwell duration | Intradialytic | All | Numeric | XXX.X | min | min | 1..1 | M | C7 | N/A |  | N/A
. . . MDCX_PDIALY_DWELL_DURATION | Dwell Duration ¦ Actual ¦ Exchange ¦ Dialysis, Peritoneal | Dwell Duration | Actual | Exchange | Dialysis, Peritoneal |  | Actual Dwell Duration | Actual dwell duration | Intradialytic | All | Numeric | XXX.X | min | min | 1..1 | M | X | N/A |  | N/A
. . . MDCX_PDIALY_DRAIN_START_TIME | Drain Start ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Drain Start | Time | Exchange | Dialysis, Peritoneal |  | Drain Start time | The time that the drain phase started. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 1..1 | C3 | X | N/A |  | N/A
. . . MDCX_PDIALY_DRAIN_END_TIME | Drain End ¦ Time ¦ Exchange ¦ Dialysis, Peritoneal | Drain End | Time | Exchange | Dialysis, Peritoneal |  | Drain End time | The time that the drain phase ended. | Intradialytic | All | Date Time | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 1..1 | C3 | X | N/A |  | N/A
. . . MDCX_PDIALY_DRAIN_MODE | Drain Mode ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Drain Mode | Setting | Exchange | Dialysis, Peritoneal |  | Drain mode | Drain mode | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 1..1 | C3 | C7 | PD_TBL_07 | TOBAG TODRAIN UNKNOWN | Cycler to Bag Cycler to drain Manual
. . . MDCX_PDIALY_DRAIN_DURATION_SETTING | Drain Duration ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Drain Duration | Setting | Exchange | Dialysis, Peritoneal |  | Prescribed Drain duration | Prescribed drain duration | Intradialytic | All | Numeric | XXX.X | min | min | 1..1 | C3 | C7 | N/A |  | N/A
. . . MDCX_PDIALY_DRAIN_DURATION | Drain Duration ¦ Actual ¦ Exchange ¦ Dialysis, Peritoneal | Drain Duration | Actual | Exchange | Dialysis, Peritoneal |  | Actual Drain duration | Actual drain duration | Intradialytic | All | Numeric | XXX.X | min | min | 1..1 | C3 | X | N/A |  | N/A
. . . MDCX_PDIALY_DRAIN_VOLUME_SETTING | Drain Volume ¦ Setting ¦ Exchange ¦ Dialysis, Peritoneal | Drain Volume | Setting | Exchange | Dialysis, Peritoneal |  | Prescribed Drain volume | Prescribed Drain volume | Intradialytic | All | Numeric | X.XXX | L | L | 1..1 | C3 | C7 | N/A |  | N/A
. . . MDCX_PDIALY_DRAIN_VOLUME | Drain Volume ¦ Actual ¦ Exchange ¦ Dialysis, Peritoneal | Drain Volume | Actual | Exchange | Dialysis, Peritoneal |  | Actual Drain volume | Actual Drain volume | Intradialytic | All | Numeric | X.XXX | L | L | 1..1 | C3 | X | N/A |  | N/A
. . MDCX_DEV_PDIALY_SUPPLY_LINE_CHAN | Subsystem ¦ Supply Line ¦ ¦ Dialysis, Peritoneal | Subsystem | Supply Line |  | Dialysis, Peritoneal |  |  |  | All | All |  |  |  |  | 0..1 | O | X
. . . MDCX_PDIALY_PRESSURE_DIALYSATE | Pressure ¦ Status ¦ Supply Line ¦ Dialysis, Peritoneal | Pressure | Status | Supply Line | Dialysis, Peritoneal |  | Dialysate Supply Pressure Sensor Value | Dialysate pressure in the supply line(s). | Intradialytic | All | Numeric | XXXX | mmHg | mm[Hg] | 0..1 | O | X | N/A |  | N/A
. . MDCX_DEV_PDIALY_DRAIN_LINE_CHAN | Subsystem ¦ Drain Line ¦ ¦ Dialysis, Peritoneal | Subsystem | Drain Line |  | Dialysis, Peritoneal |  |  |  | All | All |  |  |  |  | 0..1 | O | X
. . . MDCX_PDIALY_PRESSURE_DRAIN | Pressure ¦ Status ¦ Drain Line ¦ Dialysis, Peritoneal | Pressure | Status | Drain Line | Dialysis, Peritoneal |  | Drain Pressure Sensor Value | Dialysate pressure in the drain line. | Intradialytic | All | Numeric | XXXX | mmHg | mm[Hg] | 0..1 | O | X | N/A |  | N/A
. . MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | Subsystem ¦ Patient Line ¦ ¦ Dialysis, Peritoneal | Subsystem | Patient Line |  | Dialysis, Peritoneal |  |  |  | All | All |  |  |  |  | 0..1 | O | X
. . . MDCX_PDIALY_PRESSURE_PATIENT | Pressure ¦ Status ¦ Patient Line ¦ Dialysis, Peritoneal | Pressure | Status | Patient Line | Dialysis, Peritoneal |  | Patient Pressure Sensor Value | Dialysate pressure in the patient line. | Intradialytic | All | Numeric | XXXX | mmHg | mm[Hg] | 0..1 | O | X | N/A |  | N/A
. . MDCX_DEV_PDIALY_HEATER_LINE_CHAN | Subsystem ¦ Heater Line ¦ ¦ Dialysis, Peritoneal | Subsystem | Heater Line |  | Dialysis, Peritoneal |  |  |  | All | All |  |  |  |  | 0..1 | O | X
. . MDCX_DEV_PDIALY_MACH_STATE_CHAN | Subsystem ¦ Machine State ¦ ¦ Dialysis, Peritoneal | Subsystem | Machine State |  | Dialysis, Peritoneal |  |  |  | All | All |  |  |  |  | 0..1 | O | X
. . . MDC_VENDOR_ALERT_NUM | ErrorEvent (numeric) ¦ Safety Systen, ‹Notifications› ¦ Machine State ¦ Dialysis, Peritoneal | ErrorEvent (numeric) | Safety Systen, ‹Notifications› | Machine State | Dialysis, Peritoneal |  | Alert ID | The number associated with the current alert. If no alert is active, then this object is not reported. | All | All | Numeric | XXXX | N/A | -- | 0..1 | O | X | -- |  | -- | 0 to 9999 | 10 | 3 |  | 1 | 1 | 2953 | 2953 | 68489
. . . MDC_VENDOR_ALERT_TEXT | ErrorEvent (string) ¦ Safety Systen, ‹Notifications› ¦ Machine State ¦ Dialysis, Peritoneal | ErrorEvent (string) | Safety Systen, ‹Notifications› | Machine State | Dialysis, Peritoneal |  | Alert Text | The text name of the current alert. If no alert is active, then this object is not reported. In most case this object will contain the text shown to the user. | All | All | String | N/A | N/A | -- | 0..1 | O | X | -- |  | -- | N/A | Venous Air Detected | 1 |  | 1 | 1 | 3010 | 3010 | 68546
. . . MDCX_PDIALY_AIR_SENSOR_STATE | Air ¦ System ¦ Machine State ¦ Dialysis, Peritoneal | Air | System | Machine State | Dialysis, Peritoneal |  | Air sensor state | The state of the air sensor. | All | All | Enumeration | Alphanumeric | N/A | -- | 0..1 | O | X | PD_TBL_11 | AIR FLUID UNKNOWN | Air Fluid Unknown
. . . MDCX_PDIALY_PUMP_SPEED | Speed ¦ Pump ¦ Machine State ¦ Dialysis, Peritoneal | Speed | Pump | Machine State | Dialysis, Peritoneal |  | Pump Speed | The speed of the pump. | All | All | Numeric | XXX.X | ml/min | ml/min | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_PUMP_PRESSURE | Pressure ¦ Pump ¦ Machine State ¦ Dialysis, Peritoneal | Pressure | Pump | Machine State | Dialysis, Peritoneal |  | Pump Pressure | The pressure at the pump outlet. | All | All | Numeric | XXX.X | mmHg | mm[Hg] | 0..1 | O | X | -- |  | --
. . . MDCX_PDIALY_TEMP_DIALYSATE_IN | Temperature, In ¦ Dialysate ¦ Machine State ¦ Dialysis, Peritoneal | Temperature, In | Dialysate | Machine State | Dialysis, Peritoneal | tech high low thr | Dialysate Temperature Sensor Input | Dialysate temperature at the inlet. | All | All | Numeric | XX.X | ˚C | Cel | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_TEMP_DIALYSATE_OUT | Temperature, Out ¦ Dialysate ¦ Machine State ¦ Dialysis, Peritoneal | Temperature, Out | Dialysate | Machine State | Dialysis, Peritoneal | tech high low thr | Dialysate Temperature Sensor Output | Dialysate temperature at the outlet. | All | All | Numeric | XX.X | ˚C | Cel | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_TEMP_HEATER_BAG | Temperature ¦ Heater Bag ¦ Machine State ¦ Dialysis, Peritoneal | Temperature | Heater Bag | Machine State | Dialysis, Peritoneal |  | Dialysate Heater Bag Temperature Sensor | Dialysate temperature in the heater bag. | All | All | Numeric | XX.X | ˚C | Cel | 0..1 | O | X | N/A |  | N/A
. . . MDC_TEMP_DEVICE_INTERNAL | Temperature ¦ Device, Internal ¦ Machine State ¦ Dialysis, Peritoneal | Temperature | Device, Internal | Machine State | Dialysis, Peritoneal |  | Device Internal Temperature Monitoring | Temperature inside the cycler. | All | All | Numeric | XX.X | ˚C | Cel | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_AIR_SENSOR_IN | Air, In ¦ Patient ¦ Machine State ¦ Dialysis, Peritoneal | Air, In | Patient | Machine State | Dialysis, Peritoneal |  | Inlet Air Sensor Value | Measurement of the air going to the patient. | All | All | Numeric | XXX | mL | mL | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_AIR_SENSOR_OUT | Air, Out ¦ Patient ¦ Machine State ¦ Dialysis, Peritoneal | Air, Out | Patient | Machine State | Dialysis, Peritoneal |  | Outlet Air Sensor Value | Measurement of the air coming from the patient. | All | All | Numeric | XXX | mL | mL | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_LEAK_SENSOR | Leak ¦ ¦ Machine State ¦ Dialysis, Peritoneal | Leak |  | Machine State | Dialysis, Peritoneal |  | Leak Sensor(s) | Detects fluid leaks in the cycler. | All | All | Bool |  |  |  | 0..1 | O | X | N/A |  | N/A
. . . MDCX_ATTR_AC_MAINS_RMS | Voltage ¦ AC Mains ¦ Machine State ¦ Dialysis, Peritoneal | Voltage | AC Mains | Machine State | Dialysis, Peritoneal |  | Mains Power Monitoring | Monitors AC line in. | All | All | Numeric | XXX.X | V | V | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_VOLTAGE_MOTOR | Voltage ¦ Pump ¦ Machine State ¦ Dialysis, Peritoneal | Voltage | Pump | Machine State | Dialysis, Peritoneal |  | Motor Voltage Monitoring | Motor Voltage Monitoring | All | All | Numeric | XXX.X | V | V | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_VOLTAGE_HEATER | Voltage ¦ Heater Bag ¦ Machine State ¦ Dialysis, Peritoneal | Voltage | Heater Bag | Machine State | Dialysis, Peritoneal |  | Heater Input Voltage | The heater input voltage level is monitored. | All | All | Numeric | XXX.X | V | V | 0..1 | O | X | N/A |  | N/A
. . . MDC_ATTR_BATT_VOLTAGE | Voltage ¦ Battery ¦ Machine State ¦ Dialysis, Peritoneal | Voltage | Battery | Machine State | Dialysis, Peritoneal |  | Battery Charge Level Monitoring | The battery voltage level is monitored. | All | All | Numeric | XX.X | V | V | 0..1 | O | X | N/A |  | N/A
. . . MDCX_ATTR_POWER_SUPPLY_VOLTAGE | Voltage ¦ Power Supply ¦ Machine State ¦ Dialysis, Peritoneal | Voltage | Power Supply | Machine State | Dialysis, Peritoneal |  | Voltage Monitoring | The cycler monitors power supply voltages. | All | All | Numeric | XX.X | V | V | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_PRESSURE_COMPRESSED_AIR | Pressure ¦ Compressed Air ¦ Machine State ¦ Dialysis, Peritoneal | Pressure | Compressed Air | Machine State | Dialysis, Peritoneal |  | Pneumatic: Compressed air pressure | Air pressure is monitored. Compressed air is positive pressure. | All | All | Numeric | XXXX | mmHg | mm[Hg] | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_PRESSURE_VACUUM | Pressure ¦ Vacuum ¦ Machine State ¦ Dialysis, Peritoneal | Pressure | Vacuum | Machine State | Dialysis, Peritoneal |  | Pneumatic: Vacuum pressure | Vacuum level is monitored. Vacuum is negative pressure. | All | All | Numeric | XXXX | mmHg | mm[Hg] | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_PRESSURE_PNEUMATIC | Pressure ¦ Pneumatics ¦ Machine State ¦ Dialysis, Peritoneal | Pressure | Pneumatics | Machine State | Dialysis, Peritoneal |  | Hydraulic Fluid Pressure | Hydraulic pump fluid pressure is monitored. | All | All | Numeric | XXXX | mmHg | mm[Hg] | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_VOLUME_SUPPLY_BAG | Volume ¦ Supply Bag ¦ Machine State ¦ Dialysis, Peritoneal | Volume | Supply Bag | Machine State | Dialysis, Peritoneal |  | Dialysate Supply Bag Volume Status Monitoring | The cycler monitors the volume of solution pumped out of the supply bag(s). | All | All | Numeric | X.XXX | L | L | 0..1 | O | X | N/A
. . . MDCX_PDIALY_VOLUME_HEATER_BAG | Volume ¦ Heater Bag ¦ Machine State ¦ Dialysis, Peritoneal | Volume | Heater Bag | Machine State | Dialysis, Peritoneal |  | Dialysate Heater Bag Volume Status Monitoring | The cycler monitors the volume of solution pumped into and out of the heater bag. | All | All | Numeric | X.XXX | L | L | 0..1 | O | X | N/A |  | N/A
. . . MDCX_PDIALY_WEIGHT_HEATER_BAG | Weight ¦ Heater Bag ¦ Machine State ¦ Dialysis, Peritoneal | Weight | Heater Bag | Machine State | Dialysis, Peritoneal |  | Heater Bag Weight | Weight of the heater bag | All | All | Numeric | XXXX.X | g | g | 0..1 | O | X | -- |  | --
. MDCX_DEV_PDIALY_PATIENT_VMD | Device ¦ Patient ¦ ¦ Dialysis, Peritoneal | Device | Patient |  | Dialysis, Peritoneal |  | Patient "Device" |  | Intradialytic | All |  |  |  |  | 0..1 | O | O
. . MDCX_DEV_PATIENT_INFO | Subsystem ¦ Patient ¦ ¦ Dialysis, Peritoneal | Subsystem | Patient |  | Dialysis, Peritoneal |  | Patient channel | A channel to hold information about the patient. | Intradialytic | All |  |  |  |  | 0..1 | O | O
. . . MDCX_PDIALY_PATIENT_TRANSPORT_TYPE | Transport Type ¦ ¦ Patient ¦ Dialysis, Peritoneal | Transport Type |  | Patient | Dialysis, Peritoneal |  | Transport Type | The rate at which small molecules cross the peritoneal membrane. | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 0..1 | O | O | PD_TBL_08 | LOW MEDLOW MEDHIGH HIGH | Low Medium Low Medium High High
. . . MDCX_PDIALY_PATIENT_CATHETER_TYPE | Catheter Type ¦ ¦ Patient ¦ Dialysis, Peritoneal | Catheter Type |  | Patient | Dialysis, Peritoneal |  | Catheter Type | Catheter Type | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 0..1 | O | O | PD_TBL_09 | STRAIGHT-TENCK OREO-ZELL TWH COILED-TENCK STRAIGHT-SWAN-TENCK COILED-SWAN-TENCK STRAIGHT-MO-SWAN COLIED-MO-SWAN MON-POP VINCENZA DO-PAOLA VALLI ASH-ADV CRUZ OTHER | Straight Tenckhoff Oreopoulos-Zellerman Toronto Western Hospital Coiled Tenckhoff Straight Swan-neck Tenckhoff Coiled Swan-neck Tenckhoff Straight Missouri Swan-neck Coiled Missouri Swan-neck Moncrief-Popovich Vincenza Do Paolo Valli (Balloon) Ash Advantage (T-Fluted) Cruz (Pail-Handle) Other/Unknown
. . . MDCX_PDIALY_PATIENT_CATHETER_DATE | Catheter Date ¦ ¦ Patient ¦ Dialysis, Peritoneal | Catheter Date |  | Patient | Dialysis, Peritoneal |  | Catheter Placement date | Catheter Placement date | Intradialytic | All | Date | YYYYMMDDHHMMSS[.SSS]+/-ZZZZ | N/A | -- | 0..1 | O | O | N/A |  | N/A
. . . MDCX_PDIALY_PAT_TARGET_WT | Weight ¦ Target ¦ Patient ¦ Dialysis, Peritoneal | Weight | Target | Patient | Dialysis, Peritoneal |  | Target Weight | The patient's target weight. This is captured for information purposes only and is not used by the dialysis machine when determining treatment settings. | Intradialytic | All | Numeric | XXX.X | Kg | kg | 0..1 | O | O | N/A |  | N/A |  | 75.0
. MDC_DEV_PRESS_BLD_NONINV_VMD | Device ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Device | Pressure, noninvasive | Blood | CVS |  | Non-invasive blood pressure | Instrument for the non-invasive measurement of blood pressure. | Intradialytic | All |  |  |  |  | 0..1 | O | O | [5 empty] | P |  | MVC (2) | 1 | 5150 | 5150 | 70686
. . MDC_DEV_PRESS_BLD_NONINV_CHAN | Subsystem ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Subsystem | Pressure, noninvasive | Blood | CVS |  |  |  | Intradialytic | All |  |  |  |  | 0..1 | O | O | [5 empty] | P |  | MVC (3) | 1 | 5151 | 5151 | 70687
. . . MDCX_DIALY_BLD_MEAS_PHASE | Measurement Phase ¦ Pressure, noninvasive ¦ Blood ¦ Dialysis, Peritoneal | Measurement Phase | Pressure, noninvasive | Blood | Dialysis, Peritoneal |  | Blood Pressure Measurement Phase | The phase of the treatment when the patient should perform a blood pressure measurement. | Intradialytic | All | String |  | N/A | -- | 0..1 | O | O | PD_TBL_10 | NONE PRE POST BOTH | None Pre Post Both | N/A | BOTH
. . . MDCX_DIALY_PUL_MEAS_PHASE | Measurement Phase ¦ Pulse, noninvasive ¦ Blood ¦ Dialysis, Peritoneal | Measurement Phase | Pulse, noninvasive | Blood | Dialysis, Peritoneal |  | Pulse Measurement Phase | The phase of the treatment when the patient should perform a pulse measurement. | Intradialytic | All | String |  | N/A | -- | 0..1 | O | O | PD_TBL_10 | NONE PRE POST BOTH | None Pre Post Both | N/A | BOTH
. . . MDC_ATTR_TIME_PD_MSMT | Duration ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Duration | Pressure, noninvasive | Blood | CVS |  | Blood Pressure Measurement Interval Setting | Time allowed to elapse between automatic blood pressure measurements. A value of zero indicates that automatic measurements are not being taken. | Intradialytic | All | Numeric | XXX | minutes | min | 0..1 | O | O |  |  | -- | 0 to 999 | 30 | P |  | 1 | 1 | 2443 | 2443 | 67979
. . . MDC_ATTR_PT_BODY_POSN | Body Position (enumeration) ¦ Pressure, noninvasive ¦ Blood ¦ CVS | Body Position (enumeration) | Pressure, noninvasive | Blood | CVS |  | Patient body position | Patient body position at the time of the {blood pressure} measurement. | Intradialytic | All | Enumeration | Alphanumeric | N/A | -- | 0..1 | O | O | _TBL_04 – Patient Position | SITTING STANDING SUPINE | SITTING STANDING SUPINE | N/A | SUPINE | 3 |  | 1 | 1 | 2599 | 2599 | 68135
. . . MDC_PRESS_BLD_NONINV_DIA | Pressure ¦ Noninvasive, Diastolic ¦ Blood ¦ CVS | Pressure | Noninvasive, Diastolic | Blood | CVS | phys high low thr | Diastolic Pressure | Minimum arterial pressure during relaxation and dilatation of the ventricles of the heart when the ventricles fill with blood | Intradialytic | All | Numeric | XXX | mmHg | mm[Hg] | 0..1 | O | X |  |  | -- | 0 to 999 | 80 | P |  | SDM (2) | 2 | 18950 | 18950 | 150022
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Noninvasive, Diastolic ¦ Blood ¦ Dialysis, Peritoneal | Entry Range | Noninvasive, Diastolic | Blood | Dialysis, Peritoneal |  | Diastolic Pressure Entry Limits | The minimum and maximum value that the user can enter for a diastolic pressure. | Intradialytic | All | Numeric Range | XXX^XXX | mmHg | mm[Hg] | 0..1 | O | O |  |  |  | 0 to 999 | 0^999
. . . MDC_PULS_RATE_NON_INV | Rate ¦ Noninvasive, Pulse ¦ Blood ¦ CVS | Rate | Noninvasive, Pulse | Blood | CVS | phys high low thr | Heart Rate | Number of heart beats per minute | Intradialytic | All | Numeric | XXX | BPM | {beats}/min | 0..1 | O | X |  |  | -- | 0 to 999 | 70 | P |  | RCE (2),, | 2 | 18474 | 18474 | 149546
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Noninvasive, Pulse ¦ Blood ¦ Dialysis, Peritoneal | Entry Range | Noninvasive, Pulse | Blood | Dialysis, Peritoneal |  | Pulse Entry Limits | The minimum and maximum value that the user can enter for a pulse measurement. | Intradialytic | All | Numeric Range | XXX^XXX | BPM | {beats}/min | 0..1 | O | O |  |  |  | 0 to 999 | 20^200
. . . MDC_PRESS_BLD_NONINV_MEAN | Pressure ¦ Noninvasive, Mean ¦ Blood ¦ CVS | Pressure | Noninvasive, Mean | Blood | CVS |  | Mean Arterial Pressure | Average pressure in a patient's arteries during one cardiac cycle | Intradialytic | All | Numeric | XXX | mmHg | mm[Hg] | 0..1 | O | X |  |  | -- | 0 to 999 | 100 | P |  | SDM (3) | 2 | 18951 | 18951 | 150023
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Noninvasive, Mean ¦ Blood ¦ Dialysis, Peritoneal | Entry Range | Noninvasive, Mean | Blood | Dialysis, Peritoneal |  | Mean Arterial Pressure Entry Limits | The minimum and maximum value that the user can enter for a mean arterial pressure. | Intradialytic | All | Numeric Range | XXX^XXX | mmHg | mm[Hg] | 0..1 | O | O |  |  |  | 0 to 999 | 0^999
. . . MDC_PRESS_BLD_NONINV_SYS | Pressure ¦ Noninvasive, Systolic ¦ Blood ¦ CVS | Pressure | Noninvasive, Systolic | Blood | CVS | phys high low thr | Systolic Pressure | Maximum arterial pressure during contraction of the left ventricle of the heart | Intradialytic | All | Numeric | XXX | mmHg | mm[Hg] | 0..1 | O | X |  |  | -- | 0 to 999 | 120 | P |  | SDM (1) | 2 | 18949 | 18949 | 150021
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Noninvasive, Systolic ¦ Blood ¦ Dialysis, Peritoneal | Entry Range | Noninvasive, Systolic | Blood | Dialysis, Peritoneal |  | Systolic Pressure Entry Limits | The minimum and maximum value that the user can enter for a systolic pressure. | Intradialytic | All | Numeric Range | XXX^XXX | mmHg | mm[Hg] | 0..1 | O | O |  |  |  | 0 to 999 | 0^999
. MDC_DEV_SPEC_PROFILE_GLUCOSE | Device ¦ Glucose ¦ Blood ¦ CVS | Device | Glucose | Blood | CVS |  | Glucose "Device" |  | Intradialytic | All |  |  |  |  | 0..1 | O | O
. . MDC_DEV_CHAN | Subsystem ¦ Glucose ¦ Blood ¦ CVS | Subsystem | Glucose | Blood | CVS |  | Glucose Channel | A channel to hold information about the patient's glucose. | Intradialytic | All |  |  |  |  | 0..1 | O | O
. . . MDCX_DIALY_PAT_GLU_MEAS_PHASE | Measurement Phase ¦ Glucose ¦ Blood ¦ Dialysis, Peritoneal | Measurement Phase | Glucose | Blood | Dialysis, Peritoneal |  | Glucose Measurement Phase | The phase of the treatment when the patient should perform a glocuse measurement. | Intradialytic | All | String |  | N/A | -- | 0..1 | O | O | PD_TBL_10 | NONE PRE POST BOTH | None Pre Post Both
. . . MDC_CONC_GLU_CAPILLARY_PLASMA | Measurement ¦ Glucose ¦ Blood ¦ Dialysis, Peritoneal | Measurement | Glucose | Blood | Dialysis, Peritoneal |  | Blood Glucose | Blood glucose at the start of treatment | Intradialytic | All | Numeric | XX.X | mmol/L | mmol/L | 0..1 | O | X | N/A |  | N/A
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Glucose ¦ Blood ¦ Dialysis, Peritoneal | Entry Range | Glucose | Blood | Dialysis, Peritoneal |  | Glucose Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XX.X^XX.X | mmol/L | mmol/L | 0..1 | O | O | N/A |  | N/A
. MDC_DEV_SPEC_PROFILE_SCALE | Device ¦ Weight ¦ Scale ¦ CVS | Device | Weight | Scale | CVS |  | Scale "Device" |  | Intradialytic | All |  |  |  |  | 0..1 | O | O
. . MDC_DEV_CHAN | Subsystem ¦ Weight ¦ Scale ¦ CVS | Subsystem | Weight | Scale | CVS |  | Weight Channel | A channel to hold information about the patient's weight. | Intradialytic | All |  |  |  |  | 0..1 | O | O
. . . MDCX_DIALY_PAT_WT_MEAS_PHASE | Measurement Phase ¦ Weight ¦ Patient ¦ Dialysis, Peritoneal | Measurement Phase | Weight | Patient | Dialysis, Peritoneal |  | Weight Measurement Phase | The phase of the treatment when the patient should perform a weight measurement. | Intradialytic | All | String |  | N/A | -- | 0..1 | O | O | PD_TBL_10 | NONE PRE POST BOTH | None Pre Post Both | N/A | PRE
. . . MDC_MASS_BODY_ACTUAL | Measurement ¦ Weight ¦ Patient ¦ Dialysis, Peritoneal | Measurement | Weight | Patient | Dialysis, Peritoneal |  | Patient Weight | Weight of patient at beginning or end of treatment | Intradialytic | All | Numeric | XXX.X | Kg | kg | 0..1 | O | X | N/A |  | N/A | 0 to 999.9 | 75.9 | 3 |  | MMM | 2 |  | 27928 | 159000
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Weight ¦ Patient ¦ Dialysis, Peritoneal | Entry Range | Weight | Patient | Dialysis, Peritoneal |  | Weight Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XXX.X^XXX.X | Kg | kg | 0..1 | O | O | N/A |  | N/A | 0 to 999.9 | 60.0
. MDC_DEV_SPEC_PROFILE_TEMP | Device ¦ Temperature ¦ Thermometer ¦ CVS | Device | Temperature | Thermometer | CVS |  | Thermometer "Device" |  | Intradialytic | All |  |  |  |  | 0..1 | O | O
. . MDC_DEV_CHAN | Subsystem ¦ Temperature ¦ Thermometer ¦ CVS | Subsystem | Temperature | Thermometer | CVS |  | Temperature Channel | A channel to hold information about the patient's temperature. | Intradialytic | All |  |  |  |  | 0..1 | O | O
. . . MDCX_DIALY_PAT_TEMP_MEAS_PHASE | Measurement Phase ¦ Temperature ¦ Patient ¦ Dialysis, Peritoneal | Measurement Phase | Temperature | Patient | Dialysis, Peritoneal |  | Temperature Measurement Phase | The phase of the treatment when the patient should perform a temperature measurement. | Intradialytic | All | String |  | N/A | -- | 0..1 | O | O | PD_TBL_10 | NONE PRE POST BOTH | None Pre Post Both | N/A | NONE
. . . MDC_TEMP_TYMP | Measurement ¦ Temperature, Tympanic ¦ Patient ¦ Dialysis | Measurement | Temperature, Tympanic | Patient | Dialysis | phys high thr | Patient Temperature | Patient's temperature when measured at the ear drum. | Intradialytic | All | Numeric | XX.X | oC | Cel | 0..1 | O | X | N/A |  | N/A | 0 to 99.9 | 39.0
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Temperature, Tympanic ¦ Patient ¦ Dialysis, Peritoneal | Entry Range | Temperature, Tympanic | Patient | Dialysis, Peritoneal |  | Temperature Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XX.X^XX.X | oC | Cel | 0..1 | O | O | N/A |  | N/A | 0 to 99.9 | 35.0^41.1
. . . MDC_TEMP_ORAL | Measurement ¦ Temperature, Oral ¦ Patient ¦ Dialysis | Measurement | Temperature, Oral | Patient | Dialysis | phys high thr | Patient Temperature | Patient's temperature when measured orally. | Intradialytic | All | Numeric | XX.X | oC | Cel | 0..1 | O | X | N/A |  | N/A | 0 to 99.9 | 39.0
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Temperature, Oral ¦ Patient ¦ Dialysis, Peritoneal | Entry Range | Temperature, Oral | Patient | Dialysis, Peritoneal |  | Temperature Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XX.X^XX.X | oC | Cel | 0..1 | O | O | N/A |  | N/A | 0 to 99.9 | 35.0^41.1
. . . MDC_TEMP_BODY | Measurement ¦ Temperature, Body ¦ Patient ¦ Dialysis | Measurement | Temperature, Body | Patient | Dialysis | phys high thr | Patient Temperature | Patient's temperature when measured at an unknown location | Intradialytic | All | Numeric | XX.X | oC | Cel | 0..1 | O | X | N/A |  | N/A | 0 to 99.9 | 39.0
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Temperature, Body ¦ Patient ¦ Dialysis, Peritoneal | Entry Range | Temperature, Body | Patient | Dialysis, Peritoneal |  | Temperature Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XX.X^XX.X | oC | Cel | 0..1 | O | O | N/A |  | N/A | 0 to 99.9 | 35.0^41.1
. . . MDC_TEMP_AXILLA | Measurement ¦ Temperature, Axillary ¦ Patient ¦ Dialysis | Measurement | Temperature, Axillary | Patient | Dialysis | phys high thr | Patient Temperature | Patient's temperature when measured at the arm pit. | Intradialytic | All | Numeric | XX.X | oC | Cel | 0..1 | O | X | N/A |  | N/A | 0 to 99.9 | 39.0
. . . . MDC_ATTR_VAL_RANGE | Entry Range ¦ Temperature, Axillary ¦ Patient ¦ Dialysis, Peritoneal | Entry Range | Temperature, Axillary | Patient | Dialysis, Peritoneal |  | Temperature Entry Limits | The minimum and maximum values for entry by the user. | Intradialytic | All | Numeric Range | XX.X^XX.X | oC | Cel | 0..1 | O | O | N/A |  | N/A | 0 to 99.9 | 35.0^41.1

## Events
Logical Source | Source | Event | SysName | SysNameBase | SysName1stDesc | SysName2ndDesc | SysName3rdDesc | Alert Type | Dialysis Common Name | Definition | Usage | gp | s | Disc | PART | pCODE10 | CODE10 | CF_CODE10
‹vendor-text› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_INTERNAL_SYSTEM_FAULT | Failure ¦ System ¦ System ¦ Dialysis, Peritoneal | Failure | System | System | Dialysis, Peritoneal | tech |  | Manufacturer specific system fault. | O
‹vendor-text› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_MSMT_OUT_OF_RANGE | Out of Range ¦ Measurement ¦ System ¦ Dialysis, Peritoneal | Out of Range | Measurement | System | Dialysis, Peritoneal | tech |  | Manufacturer specific measurement out of range. | O
‹vendor-text› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_SELFTEST_FAILURE | Failure ¦ Selftest ¦ System ¦ Dialysis, Peritoneal | Failure | Selftest | System | Dialysis, Peritoneal | tech |  | Manufacturer specific selftest failure. | O
‹vendor-text› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_ADVIS_PDIALY_USER_MESSAGE | Message, User ¦ System ¦ System ¦ Dialysis, Peritoneal | Message, User | System | System | Dialysis, Peritoneal | tech |  | Machine specific advisory | O
‹vendor-text› | MDCX_DEV_PDIALY_MACHINE_MDS | MDCX_EVT_PATIENT_ID_INVALID | Invalid ¦ Patient ID ¦ System ¦ Dialysis, Peritoneal | Invalid | Patient ID | System | Dialysis, Peritoneal | tech |  | Invalid Patient ID | O
‹vendor-text› | MDCX_DEV_PDIALY_MACHINE_MDS | MDCX_EVT_TREATMENT_ERR | General Error ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | General Error | Treatment | System | Dialysis, Peritoneal | tech |  | An error occurred during treatment | O
‹vendor-text› | MDCX_DEV_PDIALY_MACHINE_MDS | MDCX_EVT_UNRECOV_TREATMENT_ERR | Unrecoverable ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Unrecoverable | Treatment | System | Dialysis, Peritoneal | tech |  | An unrecoverable error occurred during treatment | O
MDC_ATTR_PT_WEIGHT | MDC_MASS_BODY_ACTUAL | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Advisory | Treatment | System | Dialysis, Peritoneal | tech |  | The user should check the patient weight. | O
‹drain-line› | MDCX_DEV_PDIALY_DRAIN_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_BLOCKED | Blockage ¦ System ¦ System ¦ Dialysis, Peritoneal | Blockage | System | System | Dialysis, Peritoneal | tech |  | A line is blocked; checking is required. | O
‹drain-line› | MDCX_DEV_PDIALY_DRAIN_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_BLOCKED_CHK_REQD | Blockage Check ¦ System ¦ System ¦ Dialysis, Peritoneal | Blockage Check | System | System | Dialysis, Peritoneal | tech |  | Slow flow in drain line should be checked by user. | O
‹patient-and-or-drain-line› | MDCX_DEV_PDIALY_DRAIN_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_BLOCKED_OUTFLOW | Blockage ¦ Outflow ¦ System ¦ Dialysis, Peritoneal | Blockage | Outflow | System | Dialysis, Peritoneal | tech |  | Bloced flow in drain line. | O
‹drain-line› | MDCX_DEV_PDIALY_DRAIN_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_SLOW | Slow Flow ¦ System ¦ System ¦ Dialysis, Peritoneal | Slow Flow | System | System | Dialysis, Peritoneal | tech |  | Slow flow in drain line. | O
‹drain-line› | MDCX_DEV_PDIALY_DRAIN_LINE_CHAN | MDCX_EVT_PDIALY_LINE_CHK | Check ¦ Line ¦ System ¦ Dialysis, Peritoneal | Check | Line | System | Dialysis, Peritoneal | tech |  | Check drain line. | O
‹heater-line› | MDCX_DEV_PDIALY_HEATER_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_BLOCKED | Blockage ¦ System ¦ System ¦ Dialysis, Peritoneal | Blockage | System | System | Dialysis, Peritoneal | tech |  | A line is blocked | O
‹heater-line› | MDCX_DEV_PDIALY_HEATER_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_SLOW | Slow Flow ¦ System ¦ System ¦ Dialysis, Peritoneal | Slow Flow | System | System | Dialysis, Peritoneal | tech |  | Slow flow in heater line. | O
‹heater-bag-or-line› | MDCX_DEV_PDIALY_HEATER_LINE_CHAN | MDCX_EVT_PDIALY_HEATER_BAG_POSN_ERR_OR_LINE_CLAMPED_CHK_REQD | Check ¦ Heater Bag ¦ System ¦ Dialysis, Peritoneal | Check | Heater Bag | System | Dialysis, Peritoneal | tech |  | Verify that the heater bag is on the heater tray and the heater line is not clamped | O
‹heater-line› | MDCX_DEV_PDIALY_HEATER_LINE_CHAN | MDCX_EVT_PDIALY_LINE_CHK | Check ¦ Line ¦ System ¦ Dialysis, Peritoneal | Check | Line | System | Dialysis, Peritoneal | tech |  | Check heater line. | O
‹program-execution› | MDCX_DEV_PDIALY_MACH_STATE_CHAN | MDCX_EVT_PROGRAM_SYNC_ERROR | Error ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Error | Treatment | System | Dialysis, Peritoneal | tech |  | Error during monitoring of program execution synchronisation | O
‹mds› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_AC_POWER_FAIL | Failure ¦ Power ¦ System ¦ Dialysis, Peritoneal | Failure | Power | System | Dialysis, Peritoneal | tech |  | AC Power fail (or short circuit in one of the devices connected to an auxiliary outlet) | O
‹mds› ‹mds-or-vmd› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_BATT_LO | Low Voltage ¦ Battery ¦ System ¦ Dialysis, Peritoneal | Low Voltage | Battery | System | Dialysis, Peritoneal | tech | Low Battery | {device or sensor} Battery low | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_HID_MALF | Failure ¦ HID ¦ System ¦ Dialysis, Peritoneal | Failure | HID | System | Dialysis, Peritoneal | tech |  | HID failure | O
‹mds› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_INTERNAL_SYSTEM_FAULT | Failure ¦ System ¦ System ¦ Dialysis, Peritoneal | Failure | System | System | Dialysis, Peritoneal | tech |  | Unspecified system fault. | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_MSMT_OUT_OF_RANGE | Out of Range ¦ Measurement ¦ System ¦ Dialysis, Peritoneal | Out of Range | Measurement | System | Dialysis, Peritoneal | tech |  | Unspecified measurement out of range | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_SELFTEST_FAILURE | Failure ¦ Selftest ¦ System ¦ Dialysis, Peritoneal | Failure | Selftest | System | Dialysis, Peritoneal | tech |  | Unspecified selftest failure | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_UNRECOV_ERR | Failure ¦ Unrecoverable ¦ System ¦ Dialysis, Peritoneal | Failure | Unrecoverable | System | Dialysis, Peritoneal | tech |  | Unrecoverable state machine error in communication. | O
‹mds› | MDCX_DEV_PDIALY_MACHINE_MDS | MDCX_EVT_AC_POWER_RESTORED | Power Restored ¦ System ¦ System ¦ Dialysis, Peritoneal | Power Restored | System | System | Dialysis, Peritoneal | tech |  | AC power restored | O
‹mds› | MDCX_DEV_PDIALY_MACHINE_MDS | MDCX_EVT_DEVICE_TILTED | Tilt ¦ System ¦ System ¦ Dialysis, Peritoneal | Tilt | System | System | Dialysis, Peritoneal | tech |  | Device tilted | O
‹mds› | MDCX_DEV_PDIALY_MACHINE_MDS | MDCX_EVT_EXTERNAL_MEMORY_ERROR | Fault ¦ External Memory ¦ System ¦ Dialysis, Peritoneal | Fault | External Memory | System | Dialysis, Peritoneal | tech |  | External memory device error (e.g. external memory not available, cannot communicate, cannot be read, cannot be written) | O
‹mds› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_SPEAKER_FAIL | Failure ¦ Audio ¦ System ¦ Dialysis, Peritoneal | Failure | Audio | System | Dialysis, Peritoneal | tech |  | The speaker power and signals are monitored to ensure the speaker is active. | O
<vmd> | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_COMM_FAILURE | Failure ¦ Communications ¦ System ¦ Dialysis, Peritoneal | Failure | Communications | System | Dialysis, Peritoneal | tech |  | The cycler monitors the information flow to the cloud to determine that the information has been sent. | O
<vmd> | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_PUMP_FAILURE | Failure ¦ Pump ¦ System ¦ Dialysis, Peritoneal | Failure | Pump | System | Dialysis, Peritoneal | tech |  | The cycler has determined that the pump is no longer functional. | O
<vmd> | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_SET_POSITION_WRONG | Deviation ¦ Set Position ¦ System ¦ Dialysis, Peritoneal | Deviation | Set Position | System | Dialysis, Peritoneal | tech | Disposable Set Position Sensor | The disposable set status is monitored. | O
<vmd> | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_STRAIN_GAUGE_FAILURE | Failure ¦ Strain Gauge ¦ System ¦ Dialysis, Peritoneal | Failure | Strain Gauge | System | Dialysis, Peritoneal | tech | Strain Gauge Sensor Status | The strain gauge state is monitored. | O
‹mds› | MDCX_DEV_PDIALY_MACHINE_MDS | MDC_EVT_TOUCHSCREEN_FAIL | Failure ¦ Touchscreen ¦ System ¦ Dialysis, Peritoneal | Failure | Touchscreen | System | Dialysis, Peritoneal | tech |  | One or more elements touchscreen have failed. | O
‹mds› | MDCX_DEV_PDIALY_MACHINE_MDS | MDCX_EVT_DISPLAY_FAIL | Failure ¦ Display ¦ System ¦ Dialysis, Peritoneal | Failure | Display | System | Dialysis, Peritoneal | tech |  | One or more elements of the display have failed. | O
‹patient-line› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_PDIALY_DRAIN_COMPLICATION_CHK_REQD | Complication ¦ Drain ¦ System ¦ Dialysis, Peritoneal | Complication | Drain | System | Dialysis, Peritoneal | tech |  | Drain complication encountered: check patient and drain lines | O
‹patient-line+› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_PDIALY_FILL_COMPLICATION_CHK_REQD | Complication ¦ Fill ¦ System ¦ Dialysis, Peritoneal | Complication | Fill | System | Dialysis, Peritoneal | tech |  | Fill complication encountered: check heater, patient and drain lines | O
‹patient-line› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_BLOCKED | Blockage ¦ System ¦ System ¦ Dialysis, Peritoneal | Blockage | System | System | Dialysis, Peritoneal | tech |  | Tubing system blocked during inflow to patient. | O
‹patient-line› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_BLOCKED_CHK_REQD | Blockage Check ¦ System ¦ System ¦ Dialysis, Peritoneal | Blockage Check | System | System | Dialysis, Peritoneal | tech |  | Slow flow in patient line should be checked by user. | O
‹patient+› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_BLOCKED_INFLOW | Blockage ¦ Inflow ¦ System ¦ Dialysis, Peritoneal | Blockage | Inflow | System | Dialysis, Peritoneal | tech |  | Blocked flow in patient inlet line. | O
‹patient-and-or-drain-line› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_BLOCKED_OUTFLOW | Blockage ¦ Outflow ¦ System ¦ Dialysis, Peritoneal | Blockage | Outflow | System | Dialysis, Peritoneal | tech |  | Blocked flow in patient outlet line. | O
‹patient-line› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_SLOW | Slow Flow ¦ System ¦ System ¦ Dialysis, Peritoneal | Slow Flow | System | System | Dialysis, Peritoneal | tech |  | Slow flow in patient line. | O
‹patient-line› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_SLOW_INFLOW | Slow Flow ¦ Inflow ¦ System ¦ Dialysis, Peritoneal | Slow Flow | Inflow | System | Dialysis, Peritoneal | tech |  | Slow flow in patient inlet line. | O
‹patient-line› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_SLOW_OUTFLOW | Slow Flow ¦ Outflow ¦ System ¦ Dialysis, Peritoneal | Slow Flow | Outflow | System | Dialysis, Peritoneal | tech |  | Slow flow in patient outlet line. | O
‹patient-line› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_PDIALY_LINE_ABSENT | Missing ¦ Line ¦ System ¦ Dialysis, Peritoneal | Missing | Line | System | Dialysis, Peritoneal | tech |  | Missing or disconnected patient line. | O
‹patient-line› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_PDIALY_LINE_CHK | Check ¦ Line ¦ System ¦ Dialysis, Peritoneal | Check | Line | System | Dialysis, Peritoneal | tech |  | Check patient line. | O
‹patient-line› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_STAT_PDIALY_PATIENT_DRAIN_PHASE | Status ¦ Drain ¦ System ¦ Dialysis, Peritoneal | Status | Drain | System | Dialysis, Peritoneal | tech |  | Patient drain phase | O
‹patient-line› | MDCX_DEV_PDIALY_PATIENT_LINE_CHAN | MDCX_EVT_STAT_PDIALY_PATIENT_FILL_PHASE | Status ¦ Fill ¦ System ¦ Dialysis, Peritoneal | Status | Fill | System | Dialysis, Peritoneal | tech |  | Patient fill or refill phase | O
<supply-line> | MDCX_DEV_PDIALY_FLUID_1_CHAN MDCX_DEV_PDIALY_FLUID_2_CHAN MDCX_DEV_PDIALY_FLUID_3_CHAN MDCX_DEV_PDIALY_FLUID_4_CHAN MDCX_DEV_PDIALY_FLUID_5_CHAN MDCX_DEV_PDIALY_FLUID_6_CHAN MDCX_DEV_PDIALY_SUPPLY_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_BLOCKED | Blockage ¦ System ¦ System ¦ Dialysis, Peritoneal | Blockage | System | System | Dialysis, Peritoneal | tech |  | Tubing system blocked during outflow from patient. | O
<supply-line> | MDCX_DEV_PDIALY_FLUID_1_CHAN MDCX_DEV_PDIALY_FLUID_2_CHAN MDCX_DEV_PDIALY_FLUID_3_CHAN MDCX_DEV_PDIALY_FLUID_4_CHAN MDCX_DEV_PDIALY_FLUID_5_CHAN MDCX_DEV_PDIALY_FLUID_6_CHAN MDCX_DEV_PDIALY_SUPPLY_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_BLOCKED_CHK_REQD | Blockage Check ¦ System ¦ System ¦ Dialysis, Peritoneal | Blockage Check | System | System | Dialysis, Peritoneal | tech |  | Slow flow in supply line should be checked by user. | O
‹supply-line› | MDCX_DEV_PDIALY_FLUID_1_CHAN MDCX_DEV_PDIALY_FLUID_2_CHAN MDCX_DEV_PDIALY_FLUID_3_CHAN MDCX_DEV_PDIALY_FLUID_4_CHAN MDCX_DEV_PDIALY_FLUID_5_CHAN MDCX_DEV_PDIALY_FLUID_6_CHAN MDCX_DEV_PDIALY_SUPPLY_LINE_CHAN | MDCX_EVT_PDIALY_FLOW_SLOW | Slow Flow ¦ System ¦ System ¦ Dialysis, Peritoneal | Slow Flow | System | System | Dialysis, Peritoneal | tech |  | Slow flow in supply line. | O
‹supply-or-final-line› | MDCX_DEV_PDIALY_FLUID_1_CHAN MDCX_DEV_PDIALY_FLUID_2_CHAN MDCX_DEV_PDIALY_FLUID_3_CHAN MDCX_DEV_PDIALY_FLUID_4_CHAN MDCX_DEV_PDIALY_FLUID_5_CHAN MDCX_DEV_PDIALY_FLUID_6_CHAN MDCX_DEV_PDIALY_SUPPLY_LINE_CHAN | MDCX_EVT_PDIALY_LINE_ABSENT | Missing ¦ Line ¦ System ¦ Dialysis, Peritoneal | Missing | Line | System | Dialysis, Peritoneal | tech |  | Missing or disconnected dialysate line. | O
‹supply-line› | MDCX_DEV_PDIALY_FLUID_1_CHAN MDCX_DEV_PDIALY_FLUID_2_CHAN MDCX_DEV_PDIALY_FLUID_3_CHAN MDCX_DEV_PDIALY_FLUID_4_CHAN MDCX_DEV_PDIALY_FLUID_5_CHAN MDCX_DEV_PDIALY_FLUID_6_CHAN MDCX_DEV_PDIALY_SUPPLY_LINE_CHAN | MDCX_EVT_PDIALY_LINE_CHK | Check ¦ Line ¦ System ¦ Dialysis, Peritoneal | Check | Line | System | Dialysis, Peritoneal | tech |  | Check dialysate/supply line. | O
‹supply-line› | MDCX_DEV_PDIALY_FLUID_1_CHAN MDCX_DEV_PDIALY_FLUID_2_CHAN MDCX_DEV_PDIALY_FLUID_3_CHAN MDCX_DEV_PDIALY_FLUID_4_CHAN MDCX_DEV_PDIALY_FLUID_5_CHAN MDCX_DEV_PDIALY_FLUID_6_CHAN MDCX_DEV_PDIALY_SUPPLY_LINE_CHAN | MDCX_EVT_PDIALY_SOLUTION_NOT_REQD | Fluid ¦ Unused Line ¦ System ¦ Dialysis, Peritoneal | Fluid | Unused Line | System | Dialysis, Peritoneal | tech |  | Dialysate solution detected on a line where it is not required by the prescription. | O
num_pd_therapy_time_setting | MDCX_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Advisory | Treatment | System | Dialysis, Peritoneal | tech |  | The user should check the therapy time setting. | O
num_pd_volume_UF_total_setting | MDCX_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Advisory | Treatment | System | Dialysis, Peritoneal | tech |  | The user should check the total UF setting. | O
num_pd_volume_total_setting | MDCX_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Advisory | Treatment | System | Dialysis, Peritoneal | tech |  | The user should check the total volume setting. | O
‹treatment-parameter› | MDCX_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Advisory | Treatment | System | Dialysis, Peritoneal | tech |  | The user should check an unspecified treatment setting. | O
num_pd_volume_UF_total | MDCX_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_LO | Low ¦ Metric ¦ System ¦ Dialysis, Peritoneal | Low | Metric | System | Dialysis, Peritoneal | tech |  | Total UF volume is too low. | O
num_pd_drain_volume_initial_setting | MDCX_DEV_PDIALY_TREATMENT_CHAN | MDC_EVT_LO | Low ¦ Metric ¦ System ¦ Dialysis, Peritoneal | Low | Metric | System | Dialysis, Peritoneal | tech |  | Initial drain volume is too low. | O
‹vendor-text› | MDCX_DEV_PDIALY_TREATMENT_CHAN | MDCX_EVT_PDIALY_PRESCRIPTION_ERROR | Invalid ¦ Prescription ¦ System ¦ Dialysis, Peritoneal | Invalid | Prescription | System | Dialysis, Peritoneal | tech |  | Invalid prescription received. | O
num_pd_volume_UF_total | MDCX_DEV_PDIALY_TREATMENT_CHAN | MDCX_EVT_PDIALY_ULTRAFILTRATION_NEG | Too Low ¦ UF ¦ Treatment ¦ Dialysis, Peritoneal | Too Low | UF | Treatment | Dialysis, Peritoneal | tech |  | Negative Ultrafiltration | O
num_pd_volume_UF_total | MDCX_DEV_PDIALY_TREATMENT_CHAN | MDCX_EVT_PDIALY_ULTRAFILTRATION_POS | Too High ¦ UF ¦ Treatment ¦ Dialysis, Peritoneal | Too High | UF | Treatment | Dialysis, Peritoneal | tech |  | Positive Ultrafiltration | O
‹patient-weight› ‹treatment-data› | MDCX_DEV_PDIALY_TREATMENT_CHAN | MDCX_EVT_TREATMENT_DATA_INVALID | Invalid ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Invalid | Treatment | System | Dialysis, Peritoneal | tech |  | Invalid treatment data (e.g. possibly after a power failure) | O
num_pd_patient_volume | MDCX_DEV_PDIALY_TREATMENT_CHAN | MDCX_EVT_VOLUME_DEVIATION | Deviation ¦ Volume ¦ System ¦ Dialysis, Peritoneal | Deviation | Volume | System | Dialysis, Peritoneal | tech |  | Patient volume monitoring has identified a deviation | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_COMM_ERR_EXTERNAL | Fault ¦ External Comm ¦ System ¦ Dialysis, Peritoneal | Fault | External Comm | System | Dialysis, Peritoneal | tech |  | Internal communication error. | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_COMM_ERR_INTERNAL | Fault ¦ Internal Comm ¦ System ¦ Dialysis, Peritoneal | Fault | Internal Comm | System | Dialysis, Peritoneal | tech |  | Internal communication error. | O
‹treatment-data› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_DISPLAYED_VALUES_INCONSISTENT | Inconsistency ¦ Display ¦ System ¦ Dialysis, Peritoneal | Inconsistency | Display | System | Dialysis, Peritoneal | tech |  | Inconsistent safety-relevant information displayed | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_DISPOSABLE_SET_LEAK | Leak ¦ Disposable ¦ System ¦ Dialysis, Peritoneal | Leak | Disposable | System | Dialysis, Peritoneal | tech |  | Disposable set leak | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_LOAD_NEW_SET | Advice ¦ Load Set ¦ System ¦ Dialysis, Peritoneal | Advice | Load Set | System | Dialysis, Peritoneal | tech |  | Load a new set | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_LOAD_NEW_SET_AND_BAGS | Advice ¦ Load Set And Bags ¦ System ¦ Dialysis, Peritoneal | Advice | Load Set And Bags | System | Dialysis, Peritoneal | tech |  | Load new set and bags | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_ADVIS_TREATMENT_MAY_BE_INSUFFICIENT | Check ¦ Treatment, Insufficient ¦ System ¦ Dialysis, Peritoneal | Check | Treatment, Insufficient | System | Dialysis, Peritoneal | tech |  | Potentially insufficient treatment detected | O
‹mds-or-vmd› ‹any-bag-or-line› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_AIR_DETECTED | Fault ¦ Air Detected ¦ System ¦ Dialysis, Peritoneal | Fault | Air Detected | System | Dialysis, Peritoneal | tech |  | Maximum allowable air detected in cassette or lines. | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_CASSETTE_ABSENT | Missing ¦ Cassette ¦ System ¦ Dialysis, Peritoneal | Missing | Cassette | System | Dialysis, Peritoneal | tech |  | Cassette not detected | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_CASSETTE_AIR_DETECTED | Air ¦ Cassette ¦ System ¦ Dialysis, Peritoneal | Air | Cassette | System | Dialysis, Peritoneal | tech |  | Maximum allowable air detected in cassette. | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_CASSETTE_FAILURE | Fault ¦ Cassette ¦ System ¦ Dialysis, Peritoneal | Fault | Cassette | System | Dialysis, Peritoneal | tech |  | Cassette failure | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_INITIAL_DRAIN_VOLUME_UNEXPECTED | Deviation ¦ Drain Volume ¦ System ¦ Dialysis, Peritoneal | Deviation | Drain Volume | System | Dialysis, Peritoneal | tech |  | Unexpected volume (of fluid) during initial drain | O
‹mds-or-vmd› num_pd_heater_and_drain_tray_weight | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_PRIMING_ERROR | Failure ¦ Priming ¦ System ¦ Dialysis, Peritoneal | Failure | Priming | System | Dialysis, Peritoneal | tech |  | Priming error | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_RELOAD_SET | Error ¦ Cassette Loading ¦ System ¦ Dialysis, Peritoneal | Error | Cassette Loading | System | Dialysis, Peritoneal | tech |  | Reload the disposable set | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_TREATMENT_PAUSED_TOO_LONG | Timeout ¦ Paused ¦ System ¦ Dialysis, Peritoneal | Timeout | Paused | System | Dialysis, Peritoneal | tech |  | Treatment delay | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PDIALY_TREATMENT_TIME_EXCEEDED | Timeout ¦ Treatment Time ¦ System ¦ Dialysis, Peritoneal | Timeout | Treatment Time | System | Dialysis, Peritoneal | tech |  | Treatment time exceeded | O
‹mds› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_PT_SECURITY_KEY_ABSENT | Missing ¦ Security Key ¦ System ¦ Dialysis, Peritoneal | Missing | Security Key | System | Dialysis, Peritoneal | tech |  | Patient security key or other credential is missing. | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_STAT_DOOR_NOT_CLOSED_OR_LOCKED_CORRECTLY | Invalid ¦ Door Position ¦ System ¦ Dialysis, Peritoneal | Invalid | Door Position | System | Dialysis, Peritoneal | tech |  | Door is not closed or locked correctly. | O
‹mds› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_STAT_PDIALY_HEATER_BAG_FILLING | Status ¦ Fill, Heater Bag ¦ System ¦ Dialysis, Peritoneal | Status | Fill, Heater Bag | System | Dialysis, Peritoneal | tech |  | Filling heater bag | O
‹mds-or-vmd-or-chan› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_STAT_PDIALY_PT_DISCONN | Status ¦ Disconnected ¦ System ¦ Dialysis, Peritoneal | Status | Disconnected | System | Dialysis, Peritoneal | tech |  | Patient disconnect timeout (intentional) | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_STAT_PDIALY_TREATMENT_PAUSED | Status ¦ Paused ¦ System ¦ Dialysis, Peritoneal | Status | Paused | System | Dialysis, Peritoneal | tech |  | Treatment paused (e.g. stop button pressed) | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_SW_ERR | Error ¦ Software ¦ System ¦ Dialysis, Peritoneal | Error | Software | System | Dialysis, Peritoneal | tech |  | Software error | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_TIMER_ERR | Failure ¦ Timer ¦ System ¦ Dialysis, Peritoneal | Failure | Timer | System | Dialysis, Peritoneal | tech |  | Error during the plausibility check of the time counter | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_TREATMENT_RESULTS_INCONSISTENT | Inconsistent ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Inconsistent | Treatment | System | Dialysis, Peritoneal | tech |  | Inconsistent treatment results | O
‹mds-or-vmd› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_UNRECOV_TREATMENT_DATA_ERR | Unrecoverable ¦ Treatment, Data ¦ System ¦ Dialysis, Peritoneal | Unrecoverable | Treatment, Data | System | Dialysis, Peritoneal | tech |  | An unrecoverable data error occurred during treatment | O
num_pd_drain_volume | MDCX_PDIALY_DRAIN_VOLUME | MDC_EVT_HI | High ¦ Metric ¦ System ¦ Dialysis, Peritoneal | High | Metric | System | Dialysis, Peritoneal | tech |  | Drain volume too high | O
num_pd_drain_volume | MDCX_PDIALY_DRAIN_VOLUME | MDC_EVT_LO | Low ¦ Metric ¦ System ¦ Dialysis, Peritoneal | Low | Metric | System | Dialysis, Peritoneal | tech |  | Drain volume too low | O
num_pd_volume_drain_phase | MDCX_PDIALY_DRAIN_VOLUME | MDCX_EVT_PDIALY_DRAIN_PHASE_NOT_COMPLETED | Incomplete ¦ Drain ¦ System ¦ Dialysis, Peritoneal | Incomplete | Drain | System | Dialysis, Peritoneal | tech |  | Drain phase not completed | O
num_pd_volume_day_fill | MDCX_PDIALY_FILL_VOLUME_SETTING | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Advisory | Treatment | System | Dialysis, Peritoneal | tech |  | User should check the day fill volume setting. | O
num_pd_volume_fill_except_last | MDCX_PDIALY_FILL_VOLUME_SETTING | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Advisory | Treatment | System | Dialysis, Peritoneal | tech |  | User should check the fill volume setting. | O
num_pd_volume_last_fill | MDCX_PDIALY_FILL_VOLUME_SETTING | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Advisory | Treatment | System | Dialysis, Peritoneal | tech |  | User should check the last fill volume setting. | O
num_pd_volume_night_fill | MDCX_PDIALY_FILL_VOLUME_SETTING | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Advisory | Treatment | System | Dialysis, Peritoneal | tech |  | User should check the nigh fill volume setting. | O
num_pd_volume_fill_phase num_pd_volume_refill_phase | MDCX_PDIALY_FILL_VOLUME | MDCX_EVT_PDIALY_FILL_PHASE_NOT_COMPLETED | Incomplete ¦ Fill ¦ System ¦ Dialysis, Peritoneal | Incomplete | Fill | System | Dialysis, Peritoneal | tech |  | Fill or refill phase not completed | O
num_pd_drain_pressure | MDCX_PDIALY_PRESSURE_DRAIN | MDC_EVT_MSMT_OUT_OF_RANGE | Out of Range ¦ Measurement ¦ System ¦ Dialysis, Peritoneal | Out of Range | Measurement | System | Dialysis, Peritoneal | tech |  | Drain pressure out of range | O
num_pd_pressure | MDCX_PDIALY_PRESSURE_DRAIN | MDC_EVT_SELFTEST_FAILURE | Failure ¦ Selftest ¦ System ¦ Dialysis, Peritoneal | Failure | Selftest | System | Dialysis, Peritoneal | tech |  | Drain pressure selftest failure | O
num_pd_patient_pressure | MDCX_PDIALY_PRESSURE_PATIENT | MDC_EVT_HI | High ¦ Metric ¦ System ¦ Dialysis, Peritoneal | High | Metric | System | Dialysis, Peritoneal | tech |  | Patient pressure too high | O
num_pd_patient_pressure | MDCX_PDIALY_PRESSURE_PATIENT | MDC_EVT_LO | Low ¦ Metric ¦ System ¦ Dialysis, Peritoneal | Low | Metric | System | Dialysis, Peritoneal | tech |  | Patient Pressure too low | O
num_pd_patient_pressure | MDCX_PDIALY_PRESSURE_PATIENT | MDC_EVT_MSMT_ERR | Error ¦ Measurement ¦ System ¦ Dialysis, Peritoneal | Error | Measurement | System | Dialysis, Peritoneal | tech |  | Unspecified error in patient pressure | O
‹pump-pressure-sensor› | MDCX_PDIALY_PRESSURE_PATIENT | MDC_EVT_SELFTEST_FAILURE | Failure ¦ Selftest ¦ System ¦ Dialysis, Peritoneal | Failure | Selftest | System | Dialysis, Peritoneal | tech |  | Patient pressure selftest failure | O
num_pd_patient_pressure | MDCX_PDIALY_PRESSURE_PATIENT | MDCX_EVT_HDIALY_PT_LINE_PRESS_HI_LOW_POSN_CHK_REQD | Out of Range ¦ Patient Line Pressure ¦ System ¦ Dialysis, Peritoneal | Out of Range | Patient Line Pressure | System | Dialysis, Peritoneal | tech |  | Patient line pressure too high or low; check patient position | O
num_pd_pneumatic_pressure | MDCX_PDIALY_PRESSURE_PNEUMATIC | MDC_EVT_HI | High ¦ Metric ¦ System ¦ Dialysis, Peritoneal | High | Metric | System | Dialysis, Peritoneal | tech |  | Pneumatic pressure too high | O
num_pd_pneumatic_pressure | MDCX_PDIALY_PRESSURE_PNEUMATIC | MDC_EVT_LEAK | Leak ¦ System ¦ System ¦ Dialysis, Peritoneal | Leak | System | System | Dialysis, Peritoneal | tech |  | Leakage in a gas or fluid filled system detected. | O
num_pd_pneumatic_pressure | MDCX_PDIALY_PRESSURE_PNEUMATIC | MDC_EVT_LO | Low ¦ Metric ¦ System ¦ Dialysis, Peritoneal | Low | Metric | System | Dialysis, Peritoneal | tech |  | Pneumatic pressure too low | O
num_pd_pressure | MDCX_PDIALY_PRESSURE_PNEUMATIC | MDC_EVT_SELFTEST_FAILURE | Failure ¦ Selftest ¦ System ¦ Dialysis, Peritoneal | Failure | Selftest | System | Dialysis, Peritoneal | tech |  | Pneumatic pressure selftest failure | O
num_pd_pneumatic_pressure | MDCX_PDIALY_PRESSURE_PNEUMATIC | MDCX_EVT_PDIALY_VALVE_LEAK | Leak ¦ Valve ¦ System ¦ Dialysis, Peritoneal | Leak | Valve | System | Dialysis, Peritoneal | tech |  | Valve or pressure leak | O
num_pd_vacuum_pressure | MDCX_PDIALY_PRESSURE_VACUUM | MDC_EVT_HI | High ¦ Metric ¦ System ¦ Dialysis, Peritoneal | High | Metric | System | Dialysis, Peritoneal | tech |  | Vacuum pressure too high | O
num_pd_vacuum_pressure | MDCX_PDIALY_PRESSURE_VACUUM | MDC_EVT_LEAK | Leak ¦ System ¦ System ¦ Dialysis, Peritoneal | Leak | System | System | Dialysis, Peritoneal | tech |  | Leakage in a gas or fluid filled system detected. | O
num_pd_vacuum_pressure | MDCX_PDIALY_PRESSURE_VACUUM | MDC_EVT_LO | Low ¦ Metric ¦ System ¦ Dialysis, Peritoneal | Low | Metric | System | Dialysis, Peritoneal | tech |  | Vacuum pressure too low | O
num_pd_vacuum_pressure | MDCX_PDIALY_PRESSURE_VACUUM | MDC_EVT_MSMT_OUT_OF_RANGE | Out of Range ¦ Measurement ¦ System ¦ Dialysis, Peritoneal | Out of Range | Measurement | System | Dialysis, Peritoneal | tech |  | Vacuum pressure out of range | O
num_pd_pressure | MDCX_PDIALY_PRESSURE_VACUUM | MDC_EVT_SELFTEST_FAILURE | Failure ¦ Selftest ¦ System ¦ Dialysis, Peritoneal | Failure | Selftest | System | Dialysis, Peritoneal | tech |  | Vacuum pressure selftest failure | O
| MDCX_PDIALY_TEMP_DIALYSATE_IN | MDC_EVT_HI | High ¦ Metric ¦ Dialysate, Temp, In ¦ Dialysis, Peritoneal | High | Metric | Dialysate, Temp, In | Dialysis, Peritoneal | tech |  | Dialysate inlet temperature too high | O
| MDCX_PDIALY_TEMP_DIALYSATE_IN | MDC_EVT_LO | Low ¦ Metric ¦ Dialysate, Temp, In ¦ Dialysis, Peritoneal | Low | Metric | Dialysate, Temp, In | Dialysis, Peritoneal | tech |  | Dialysate inlet temperature too low | O
num_pd_dialysate_temperature | MDCX_PDIALY_TEMP_DIALYSATE_OUT | MDC_EVT_HI | High ¦ Metric ¦ Dialysate, Temp, Out ¦ Dialysis, Peritoneal | High | Metric | Dialysate, Temp, Out | Dialysis, Peritoneal | tech |  | Dialysate temperature too high | O
num_pd_dialysate_temperature | MDCX_PDIALY_TEMP_DIALYSATE_OUT | MDC_EVT_LO | Low ¦ Metric ¦ Dialysate, Temp, Out ¦ Dialysis, Peritoneal | Low | Metric | Dialysate, Temp, Out | Dialysis, Peritoneal | tech |  | Dialysate temperature too low | O
num_pd_dialysate_temperature | MDCX_PDIALY_TEMP_DIALYSATE_OUT | MDC_EVT_MSMT_ERR | Error ¦ Measurement ¦ Dialysate, Temp, Out ¦ Dialysis, Peritoneal | Error | Measurement | Dialysate, Temp, Out | Dialysis, Peritoneal | tech |  | Unspecified error in dialysate temperature | O
num_pd_dialysate_temperature | MDCX_PDIALY_TEMP_DIALYSATE_OUT | MDC_EVT_MSMT_OUT_OF_RANGE | Out of Range ¦ Measurement ¦ Dialysate, Temp, Out ¦ Dialysis, Peritoneal | Out of Range | Measurement | Dialysate, Temp, Out | Dialysis, Peritoneal | tech |  | Dialysate temperature out of range | O
num_pd_dialysate_temperature | MDCX_PDIALY_TEMP_DIALYSATE_OUT | MDCX_EVT_STAT_SOLUTION_WARMING | Status ¦ Warming ¦ Dialysate, Temp, Out ¦ Dialysis, Peritoneal | Status | Warming | Dialysate, Temp, Out | Dialysis, Peritoneal | tech |  | Dialysate solution is warming | O
num_pd_heater_temperature | MDCX_PDIALY_TEMP_HEATER_BAG | MDC_EVT_HI | High ¦ Metric ¦ System ¦ Dialysis, Peritoneal | High | Metric | System | Dialysis, Peritoneal | tech |  | Heater bag temperature too high | O
num_pd_heater_temperature | MDCX_PDIALY_TEMP_HEATER_BAG | MDC_EVT_LO | Low ¦ Metric ¦ System ¦ Dialysis, Peritoneal | Low | Metric | System | Dialysis, Peritoneal | tech |  | Heater bag temperature too low | O
num_pd_temperature | MDCX_PDIALY_TEMP_HEATER_BAG | MDC_EVT_MSMT_OUT_OF_RANGE | Out of Range ¦ Measurement ¦ System ¦ Dialysis, Peritoneal | Out of Range | Measurement | System | Dialysis, Peritoneal | tech |  | Heater Bag Temperature out of range | O
num_pd_heater_temperature | MDCX_PDIALY_TEMP_HEATER_BAG | MDC_EVT_MSMT_OUT_OF_RANGE_HI | High ¦ Measurement ¦ System ¦ Dialysis, Peritoneal | High | Measurement | System | Dialysis, Peritoneal | tech |  | Heater bag temperature too high | O
num_pd_heater_temperature | MDCX_PDIALY_TEMP_HEATER_BAG | MDC_EVT_SELFTEST_FAILURE | Failure ¦ Selftest ¦ System ¦ Dialysis, Peritoneal | Failure | Selftest | System | Dialysis, Peritoneal | tech |  | Heater bag temperature selftest failure | O
num_pd_heater_temperature | MDCX_PDIALY_TEMP_HEATER_BAG | MDCX_EVT_TEMP_ERR | Error ¦ Temperature ¦ System ¦ Dialysis, Peritoneal | Error | Temperature | System | Dialysis, Peritoneal | tech |  | Temperature sensor error | O
num_pd_volume_tidal_setting | MDCX_PDIALY_TIDAL_PERCENTAGE | MDC_EVT_ADVIS_SETTING_CHK | Advisory ¦ Treatment ¦ System ¦ Dialysis, Peritoneal | Advisory | Treatment | System | Dialysis, Peritoneal | tech |  | The user should check the tidal setting. | O
‹any-valve› | MDCX_DEV_PDIALY_VMD | MDCX_EVT_VALVE_MALF | Failure ¦ Valve ¦ System ¦ Dialysis, Peritoneal | Failure | Valve | System | Dialysis, Peritoneal | tech |  | Valve malfunction | O
num_pd_heater_and_drain_tray_weight | MDCX_PDIALY_WEIGHT_HEATER_BAG | MDC_EVT_HI_GT_LIM | Out of Range ¦ Metric ¦ System ¦ Dialysis, Peritoneal | Out of Range | Metric | System | Dialysis, Peritoneal | tech |  | A metric exceeds a given threshold | O
num_pd_heater_and_drain_tray_weight | MDCX_PDIALY_WEIGHT_HEATER_BAG | MDC_EVT_MSMT_ERR | Error ¦ Measurement ¦ System ¦ Dialysis, Peritoneal | Error | Measurement | System | Dialysis, Peritoneal | tech |  | Unspecified error in heater bag weight | O
num_pd_heater_and_drain_tray_weight | MDCX_PDIALY_WEIGHT_HEATER_BAG | MDC_EVT_MSMT_FAIL | Failure ¦ Measurement ¦ System ¦ Dialysis, Peritoneal | Failure | Measurement | System | Dialysis, Peritoneal | tech |  | Failure in heater bag weight | O
num_pd_heater_and_drain_tray_weight | MDCX_PDIALY_WEIGHT_HEATER_BAG | MDCX_EVT_PDIALY_ABRUPT_WEIGHT_CHANGE | Abrupt Change ¦ Weight ¦ System ¦ Dialysis, Peritoneal | Abrupt Change | Weight | System | Dialysis, Peritoneal | tech |  | Abrupt weight change; possible scale interference. | O
num_pd_heater_bag_weight num_pd_heater_and_drain_tray_weight | MDCX_PDIALY_WEIGHT_HEATER_BAG | MDCX_EVT_PDIALY_FLUID_BAG_WEIGHT_VOLUME_ERROR | Inconsistent ¦ Heater Bag Weight ¦ System ¦ Dialysis, Peritoneal | Inconsistent | Heater Bag Weight | System | Dialysis, Peritoneal | tech |  | Fluid bag weight scale (volume) measurement error. | O
num_pd_heater_bag_weight | MDCX_PDIALY_WEIGHT_HEATER_BAG | MDCX_EVT_PDIALY_HEATER_BAG_ON_HEATER_TRAY_UNEXPECTED | Deviation ¦ Heater Bag ¦ System ¦ Dialysis, Peritoneal | Deviation | Heater Bag | System | Dialysis, Peritoneal | tech |  | Heater bag detected on the heater tray when not expected. | O
num_pd_heater_and_drain_tray_weight | MDCX_PDIALY_WEIGHT_HEATER_BAG | MDCX_EVT_WEIGHT_INVALID_CHK_REQD | Invalid ¦ Measurement ¦ System ¦ Dialysis, Peritoneal | Invalid | Measurement | System | Dialysis, Peritoneal | tech |  | The dialysis machine has detected an invalid weight change and the user must verify/confirm the weight. | O
| MDC_PRESS_BLD_NONINV_DIA | MDC_EVT_HI | High ¦ Metric ¦ Diastolic ¦ NIBP | High | Metric | Diastolic | NIBP | phys | Diastolic High | Diastolic pressure is too high. | O
| MDC_PRESS_BLD_NONINV_DIA | MDC_EVT_LO | Low ¦ Metric ¦ Diastolic ¦ NIBP | Low | Metric | Diastolic | NIBP | phys | Diastolic Low | Diastolic pressure is too low. | O
| MDC_PRESS_BLD_NONINV_SYS | MDC_EVT_HI | High ¦ Metric ¦ Systolic ¦ NIBP | High | Metric | Systolic | NIBP | phys | Systolic High | Systolic pressure is too high. | O
| MDC_PRESS_BLD_NONINV_SYS | MDC_EVT_LO | Low ¦ Metric ¦ Systolic ¦ NIBP | Low | Metric | Systolic | NIBP | phys | Systolic Low | Systolic pressure is too low. | O
| MDC_PULS_RATE_NON_INV | MDC_EVT_HI | High ¦ Metric ¦ Pulse ¦ NIBP | High | Metric | Pulse | NIBP | phys | Pulse High | Pulse is too high. | O
| MDC_PULS_RATE_NON_INV | MDC_EVT_LO | Low ¦ Metric ¦ Pulse ¦ NIBP | Low | Metric | Pulse | NIBP | phys | Pulse Low | Pulse is too low. | O
| MDC_TEMP_TYMP | MDC_EVT_HI | High ¦ Metric ¦ Temperature, Tympanic ¦ Thermometer | High | Metric | Temperature, Tympanic | Thermometer | phys | Temperature High | Temperature is too high | O
| MDC_TEMP_ORAL | MDC_EVT_HI | High ¦ Metric ¦ Temperature, Oral ¦ Thermometer | High | Metric | Temperature, Oral | Thermometer | phys | Temperature High | Temperature is too high | O
| MDC_TEMP_BODY | MDC_EVT_HI | High ¦ Metric ¦ Temperature, Body ¦ Thermometer | High | Metric | Temperature, Body | Thermometer | phys | Temperature High | Temperature is too high | O
| MDC_TEMP_AXILLA | MDC_EVT_HI | High ¦ Metric ¦ Temperature, Axilla ¦ Thermometer | High | Metric | Temperature, Axilla | Thermometer | phys | Temperature High | Temperature is too high | O

## ConditionCodes
Code | Condition | Expression
C1 | The Exchange Phase is Fill or greater | MDCX_PDIALY_CURRENT_PHASE = FILL or DWELL or DRAIN or COMPLETE
C2 | The Exchange Phase is Dwell or greater | MDCX_PDIALY_CURRENT_PHASE = DWELL or DRAIN or COMPLETE
C3 | The Exchange Phase is Drain or greater | MDCX_PDIALY_CURRENT_PHASE = DRAIN or COMPLETE
C4 | The Fluid Source is bags. | MDCX_PDIALY_FLUID_SOURCE = BAG
C5 | There are APD exchanges | MDCX_PDIALY_TREAT_TYPE = APD
C6 | There are CAPD exchanges | MDCX_PDIALY_TREAT_TYPE = CAPD
C7 | The prescription type is exchange based. | MDCX_PDIALY_PRESCRIPTION_TYPE = EXCH